## Initial data
Instrumental battery to be tested: gait platform, interphase and P.C.):
CALIBRATION MATRIX
| x1 | x2 | x3 | x4 | x5 | x6 | x7 | x8
y1 | 1097.9 | 15.6 | -7.3 | -7.5 | 2.4 | 11.9
y2 | 3.5 | 1196.1 | 1.4 | -4.7 | -6.9 | -17
y3 | 42.1 | -20.8 | 1894.6 | 8.4 | -6.5 | -18.1
y4 | -0.5 | -81.4 | 1.1 | 597.3 | -1.7 | -1.7
y5 | 70.7 | 0.6 | 1 | -1.8 | 418.3 | 1.4
y6 | -0.1 | -3.3 | 1.3 | 1.2 | 2.2 | 276.7
y7
y8
SAMPLE FREQUENCY | x4: 1000; x5: Hz
REPETITION TIME | x4: 60; x5: s
Variables a registrar
VERTICAL FORCE (z)
ANT-POS FORCE (x)
MED-LAT FORCE (y)
FILTER | bar: x2
Pasa | bar: x2
CUT FREQUENCY | x2: 1 Hz
CALIBRATION WEIGHTS
# | MASS (kg) | CODE | Identification | certificate # | CONVENTIONAL MASS | UNCERTAINTY (mg) | OFFSET (g)
1 | 20 | M-015 | M-015 | MC-CC-11-05-032F | 19.9996 | 75 | -
2 | 20 | M-016 | M-016 | MC-CC-11-05-035F | 19.9992 | 75 | -
3 | 20 | M-022 | M-022 | MC-CC-11-05-031F | 19.9996 | 75 | -
4 | 20 | M-018 | M-018 | MC-CC-11-05-029F | 19.9993 | 76 | -
5 | 20 | M-019 | M-019 | MC-CC-11-05-033F | 19.9993 | 75 | -
6 | 20 | M-020 | M-020 | MC-CC-11-05-026F | 19.9989 | 75 | -
7 | 20 | M-021 | M-021 | MC-CC-11-05-034F | 19.9992 | 74 | -
8 | 10 | M-007 | M-007 | MC-P-197-2010 | 10.0002 | 58 | -1
9 | 5 | M-005 | M-005 | MC-P-266-2010 | 5 | 58 | -0.7
10 | 5 | M-006 | M-006 | MC-P-267-2010 | 5.0001 | 58 | -0.7
REFERENCE VALUE ADJUSTED | x2: 0; x3: kg

## STABILITY TESTS
UTILIZED WEIGHTS |  | Código | Código |  |  | 1 | Gravity (m/s^2)
LEFT FOOT |  | M-015
RIGHT FOOT |  | M-016
A 0 kg -- REP I | [7 empty] | A 40 kg -- REP I
Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg) |  | Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg)
0-1 | 1.0686 | 0.53141 | 0.53169 | -0.228 | -0.22787 | 0.074701 |  | 0-1 | 1.048 | 0.59088 | 0.53261 | 392.7289 | -1.1598 | -1.3648
1-2 | 1.0361 | 0.52283 | 0.53178 | -0.22444 | -0.069029 | 0.086819 |  | 1-2 | 1.0832 | 0.52107 | 0.52704 | 392.7362 | -1.1555 | -1.3726
2-3 | 1.0586 | 0.71147 | 0.55274 | -0.26255 | 0.13743 | 0.11581 |  | 2-3 | 1.0214 | 0.59348 | 0.53867 | 392.799 | -1.2219 | -1.3748
3-4 | 1.0829 | 0.52305 | 0.53622 | -0.2027 | -0.097927 | 0.13405 |  | 3-4 | 1.05 | 0.58802 | 0.5481 | 392.7867 | -1.3029 | -1.3706
4-5 | 1.0139 | 0.69256 | 0.55881 | -0.18264 | 0.039095 | 0.14067 |  | 4-5 | 1.0752 | 0.5292 | 0.51889 | 392.8204 | -1.2329 | -1.4014
5-10 | 1.0478 | 0.62896 | 0.54539 | -0.14222 | 0.11572 | 0.10824 |  | 5-10 | 1.0568 | 0.55514 | 0.54205 | 392.8319 | -1.2193 | -1.3894
10-15 | 1.0575 | 0.55485 | 0.53969 | -0.22027 | -0.019128 | 0.13042 |  | 10-15 | 1.0453 | 0.59534 | 0.5352 | 392.9508 | -1.1982 | -1.4054
15-20 | 1.0566 | 0.56263 | 0.53905 | -0.23664 | 0.010816 | 0.11865 |  | 15-20 | 1.0587 | 0.53739 | 0.5415 | 393.016 | -1.0233 | -1.4094
20-25 | 1.0636 | 0.59855 | 0.53836 | -0.17679 | -0.096883 | 0.083046 |  | 20-25 | 1.0522 | 0.5638 | 0.53323 | 393.0229 | -1.1992 | -1.405
25-30 | 1.0594 | 0.59366 | 0.54467 | -0.2046 | -0.12948 | 0.080226 |  | 25-30 | 1.0368 | 0.5599 | 0.53864 | 393.1108 | -1.0791 | -1.4033
30-35 | 1.0643 | 0.57852 | 0.54983 | -0.22472 | -0.14886 | 0.091318 |  | 30-35 | 1.0509 | 0.54838 | 0.52867 | 393.1777 | -1.1775 | -1.394
35-40 | 1.0703 | 0.59549 | 0.54109 | -0.25378 | -0.14032 | 0.083499 |  | 35-40 | 1.0537 | 0.58076 | 0.53543 | 393.1676 | -1.2481 | -1.4272
40-45 | 1.0599 | 0.56766 | 0.53342 | -0.25085 | -0.10535 | 0.086418 |  | 40-45 | 1.0451 | 0.58682 | 0.53892 | 393.272 | -1.1496 | -1.4227
45-50 | 1.0486 | 0.54312 | 0.53094 | -0.32387 | -0.28087 | 0.068213 |  | 45-50 | 1.045 | 0.59075 | 0.54637 | 393.2448 | -1.1396 | -1.4312
50-55 | 1.0523 | 0.60125 | 0.53713 | -0.24946 | -0.1302 | 0.085329 |  | 50-55 | 1.0491 | 0.53366 | 0.53365 | 393.3224 | -1.1293 | -1.4537
55-60 | 1.0517 | 0.62458 | 0.53123 | -0.26949 | 0.19917 | 0.07551 |  | 55-60 | 1.0616 | 0.55667 | 0.54095 | 393.3687 | -1.2993 | -1.4442
A 0 kg -- REP II | [7 empty] | A 40 kg -- REP II
Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg) |  | Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg)
0-1 | 1.0867 | 0.5256 | 0.54204 | 0.092255 | 0.026859 | 0.0035815 |  | 0-1 | 1.0831 | 0.53185 | 0.56339 | 391.5262 | -0.70543 | -1.1877
1-2 | 1.0817 | 0.52345 | 0.54926 | 0.059538 | 0.04801 | 0.040805 |  | 1-2 | 1.1029 | 0.54433 | 0.53103 | 391.575 | -1.143 | -1.2083
2-3 | 1.099 | 0.60005 | 0.55963 | 0.10423 | -0.057706 | 0.017514 |  | 2-3 | 1.0838 | 0.5101 | 0.54737 | 391.5147 | -0.81745 | -1.2062
3-4 | 1.0741 | 0.58162 | 0.56827 | 0.074091 | 0.13435 | 0.036474 |  | 3-4 | 1.1241 | 0.63429 | 0.55237 | 391.5653 | -0.8247 | -1.2082
4-5 | 1.0901 | 0.53558 | 0.53806 | 0.15212 | 0.07082 | 0.063863 |  | 4-5 | 1.0732 | 0.52933 | 0.54365 | 391.6167 | -1.1522 | -1.2044
5-10 | 1.0726 | 0.58662 | 0.54336 | 0.080383 | 0.029855 | 0.042557 |  | 5-10 | 1.0846 | 0.57513 | 0.5397 | 391.6056 | -1.0365 | -1.188
10-15 | 1.0704 | 0.59448 | 0.55129 | 0.058779 | 0.02189 | 0.012053 |  | 10-15 | 1.0515 | 0.5509 | 0.54433 | 391.7393 | -1.0268 | -1.2059
15-20 | 1.0767 | 0.59516 | 0.54867 | 0.067816 | 0.032606 | 0.0099832 |  | 15-20 | 1.0662 | 0.57425 | 0.53904 | 391.8189 | -0.92956 | -1.2077
20-25 | 1.0771 | 0.58459 | 0.54799 | 0.0037577 | 0.1044 | 0.015962 |  | 20-25 | 1.0727 | 0.54763 | 0.54149 | 391.8433 | -0.85809 | -1.2042
25-30 | 1.0803 | 0.60756 | 0.55772 | -0.015408 | 0.1145 | 0.0089466 |  | 25-30 | 1.0838 | 0.58131 | 0.54457 | 391.8353 | -1.0357 | -1.2182
30-35 | 1.0797 | 0.58664 | 0.54554 | -0.026153 | -0.0048416 | -0.0012707 |  | 30-35 | 1.0849 | 0.56601 | 0.54376 | 391.8895 | -0.84545 | -1.1971
35-40 | 1.0843 | 0.55341 | 0.54025 | 0.083424 | 0.04094 | -0.019099 |  | 35-40 | 1.0679 | 0.5744 | 0.53805 | 391.8731 | -0.95359 | -1.2051
40-45 | 1.08 | 0.6031 | 0.53815 | 0.036605 | -0.0063662 | -0.017192 |  | 40-45 | 1.062 | 0.53497 | 0.54868 | 391.8712 | -0.99073 | -1.2281
45-50 | 1.0734 | 0.59979 | 0.54435 | 0.068616 | -0.062039 | 0.009048 |  | 45-50 | 1.0847 | 0.5723 | 0.54739 | 391.9052 | -0.9495 | -1.2071
50-55 | 1.0951 | 0.54407 | 0.54262 | 0.1113 | -0.1063 | 0.0071234 |  | 50-55 | 1.0802 | 0.57311 | 0.548 | 391.8995 | -1.177 | -1.2126
55-60 | 1.0769 | 0.57469 | 0.54715 | 0.10181 | 0.030497 | 0.015581 |  | 55-60 | 1.0821 | 0.56922 | 0.54213 | 391.9487 | -0.99922 | -1.1863
A 0 kg -- REP III | [7 empty] | A 40 kg -- REP III
Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg) |  | Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg)
0-1 | 1.0856 | 0.51031 | 0.54737 | -0.047251 | -0.1652 | -0.00053852 |  | 0-1 | 1.0475 | 0.53426 | 0.54931 | 391.8721 | -1.2058 | -1.294
1-2 | 1.0869 | 0.54316 | 0.54531 | 0.0033396 | 0.46231 | 0.031009 |  | 1-2 | 1.09 | 0.56048 | 0.54887 | 391.8927 | -1.3475 | -1.2783
2-3 | 1.0896 | 0.56666 | 0.54821 | -0.04101 | 0.019387 | 0.027204 |  | 2-3 | 1.1165 | 0.60027 | 0.53572 | 391.9452 | -1.3769 | -1.3025
3-4 | 1.0931 | 0.50415 | 0.54783 | -0.0020059 | -0.080069 | 0.011984 |  | 3-4 | 1.1002 | 0.58754 | 0.53595 | 391.985 | -1.1689 | -1.3281
4-5 | 1.0908 | 0.50973 | 0.53083 | 0.0035428 | -0.077742 | 0.077492 |  | 4-5 | 1.0705 | 0.56395 | 0.55607 | 392.0212 | -1.3653 | -1.3042
5-10 | 1.0862 | 0.58288 | 0.54077 | -0.016408 | 0.18522 | 0.020909 |  | 5-10 | 1.0877 | 0.55115 | 0.53483 | 392.109 | -1.4061 | -1.3106
10-15 | 1.0858 | 0.57699 | 0.55131 | -0.12518 | 0.044114 | 0.037324 |  | 10-15 | 1.0666 | 0.58608 | 0.54346 | 392.2801 | -1.3515 | -1.3085
15-20 | 1.0856 | 0.5524 | 0.53856 | -0.086597 | 0.052389 | 0.032576 |  | 15-20 | 1.1038 | 0.54268 | 0.55059 | 392.3192 | -1.2825 | -1.2906
20-25 | 1.0905 | 0.60414 | 0.54955 | -0.046282 | 0.12678 | 0.051539 |  | 20-25 | 1.1035 | 0.6448 | 0.5528 | 392.4556 | -1.5156 | -1.3231
25-30 | 1.0899 | 0.57805 | 0.54771 | -0.037703 | 0.1758 | 0.027972 |  | 25-30 | 1.0991 | 0.58305 | 0.55697 | 392.5556 | -1.3305 | -1.3069
30-35 | 1.0917 | 0.63746 | 0.54858 | -0.0075077 | 0.21176 | 0.022856 |  | 30-35 | 1.0828 | 0.5741 | 0.54615 | 392.5555 | -1.4654 | -1.3134
35-40 | 1.0801 | 0.58507 | 0.54424 | -0.064935 | 0.15795 | 0.038406 |  | 35-40 | 1.0904 | 0.59666 | 0.55636 | 392.603 | -1.3588 | -1.2898
40-45 | 1.0873 | 0.63221 | 0.55452 | -0.15361 | 0.064504 | 0.024596 |  | 40-45 | 1.0786 | 0.57708 | 0.55112 | 392.5921 | -1.4087 | -1.304
45-50 | 1.0835 | 0.59754 | 0.54613 | -0.19108 | 0.042561 | 0.014178 |  | 45-50 | 1.0988 | 0.56621 | 0.54799 | 392.6519 | -1.4396 | -1.2887
50-55 | 1.0804 | 0.63024 | 0.54627 | -0.16184 | 0.17562 | 0.040374 |  | 50-55 | 1.0868 | 0.57754 | 0.54995 | 392.6491 | -1.4012 | -1.3074
55-60 | 1.0737 | 0.57593 | 0.55429 | -0.20662 | 0.038364 | 0.020724 |  | 55-60 | 1.0785 | 0.5904 | 0.54909 | 392.7024 | -1.3624 | -1.2932
MAX -- 0 kg | [7 empty] | MAX -- 40 kg
Intervalo (s) | D.E.Fzprom (kg) | D.E.Fyprom (kg) | D.E.Fxprom (kg) | Fzprom (kg) | Fyprom (kg) | Fxprom (kg) |  | Intervalo (s) | D.E.Fzprom (kg) | D.E.Fyprom (kg) | D.E.Fxprom (kg) | Fzprom (kg) | Fyprom (kg) | Fxprom (kg)
0-1 | 1.0867 | 0.53141 | 0.54737 | 0.092255 | 0.026859 | 0.074701 |  | 0-1 | 1.0831 | 0.59088 | 0.56339 | 392.7289 | -0.70543 | -1.1877
1-2 | 1.0869 | 0.54316 | 0.54926 | 0.059538 | 0.46231 | 0.086819 |  | 1-2 | 1.1029 | 0.56048 | 0.54887 | 392.7362 | -1.143 | -1.2083
2-3 | 1.099 | 0.71147 | 0.55963 | 0.10423 | 0.13743 | 0.11581 |  | 2-3 | 1.1165 | 0.60027 | 0.54737 | 392.799 | -0.81745 | -1.2062
3-4 | 1.0931 | 0.58162 | 0.56827 | 0.074091 | 0.13435 | 0.13405 |  | 3-4 | 1.1241 | 0.63429 | 0.55237 | 392.7867 | -0.8247 | -1.2082
4-5 | 1.0908 | 0.69256 | 0.55881 | 0.15212 | 0.07082 | 0.14067 |  | 4-5 | 1.0752 | 0.56395 | 0.55607 | 392.8204 | -1.1522 | -1.2044
5-10 | 1.0862 | 0.62896 | 0.54539 | 0.080383 | 0.18522 | 0.10824 |  | 5-10 | 1.0877 | 0.57513 | 0.54205 | 392.8319 | -1.0365 | -1.188
10-15 | 1.0858 | 0.59448 | 0.55131 | 0.058779 | 0.044114 | 0.13042 |  | 10-15 | 1.0666 | 0.59534 | 0.54433 | 392.9508 | -1.0268 | -1.2059
15-20 | 1.0856 | 0.59516 | 0.54867 | 0.067816 | 0.052389 | 0.11865 |  | 15-20 | 1.1038 | 0.57425 | 0.55059 | 393.016 | -0.92956 | -1.2077
20-25 | 1.0905 | 0.60414 | 0.54955 | 0.0037577 | 0.12678 | 0.083046 |  | 20-25 | 1.1035 | 0.6448 | 0.5528 | 393.0229 | -0.85809 | -1.2042
25-30 | 1.0899 | 0.60756 | 0.55772 | -0.015408 | 0.1758 | 0.080226 |  | 25-30 | 1.0991 | 0.58305 | 0.55697 | 393.1108 | -1.0357 | -1.2182
30-35 | 1.0917 | 0.63746 | 0.54983 | -0.0075077 | 0.21176 | 0.091318 |  | 30-35 | 1.0849 | 0.5741 | 0.54615 | 393.1777 | -0.84545 | -1.1971
35-40 | 1.0843 | 0.59549 | 0.54424 | 0.083424 | 0.15795 | 0.083499 |  | 35-40 | 1.0904 | 0.59666 | 0.55636 | 393.1676 | -0.95359 | -1.2051
40-45 | 1.0873 | 0.63221 | 0.55452 | 0.036605 | 0.064504 | 0.086418 |  | 40-45 | 1.0786 | 0.58682 | 0.55112 | 393.272 | -0.99073 | -1.2281
45-50 | 1.0835 | 0.59979 | 0.54613 | 0.068616 | 0.042561 | 0.068213 |  | 45-50 | 1.0988 | 0.59075 | 0.54799 | 393.2448 | -0.9495 | -1.2071
50-55 | 1.0951 | 0.63024 | 0.54627 | 0.1113 | 0.17562 | 0.085329 |  | 50-55 | 1.0868 | 0.57754 | 0.54995 | 393.3224 | -1.1293 | -1.2126
55-60 | 1.0769 | 0.62458 | 0.55429 | 0.10181 | 0.19917 | 0.07551 |  | 55-60 | 1.0821 | 0.5904 | 0.54909 | 393.3687 | -0.99922 | -1.1863
MIN -- 0 kg | [7 empty] | MIN -- 40 kg
Intervalo (s) | D.E.Fzprom (kg) | D.E.Fyprom (kg) | D.E.Fxprom (kg) | Fzprom (kg) | Fyprom (kg) | Fxprom (kg) |  | Intervalo (s) | D.E.Fzprom (kg) | D.E.Fyprom (kg) | D.E.Fxprom (kg) | Fzprom (kg) | Fyprom (kg) | Fxprom (kg)
0-1 | 1.0856 | 0.51031 | 0.54204 | -0.047251 | -0.1652 | -0.00053852 |  | 0-1 | 1.0475 | 0.53185 | 0.54931 | 391.5262 | -1.2058 | -1.294
1-2 | 1.0817 | 0.52345 | 0.54531 | 0.0033396 | 0.04801 | 0.031009 |  | 1-2 | 1.09 | 0.54433 | 0.53103 | 391.575 | -1.3475 | -1.2783
2-3 | 1.0896 | 0.56666 | 0.54821 | -0.04101 | -0.057706 | 0.017514 |  | 2-3 | 1.0838 | 0.5101 | 0.53572 | 391.5147 | -1.3769 | -1.3025
3-4 | 1.0741 | 0.50415 | 0.54783 | -0.0020059 | -0.080069 | 0.011984 |  | 3-4 | 1.1002 | 0.58754 | 0.53595 | 391.5653 | -1.1689 | -1.3281
4-5 | 1.0901 | 0.50973 | 0.53083 | 0.0035428 | -0.077742 | 0.063863 |  | 4-5 | 1.0705 | 0.52933 | 0.54365 | 391.6167 | -1.3653 | -1.3042
5-10 | 1.0726 | 0.58288 | 0.54077 | -0.016408 | 0.029855 | 0.020909 |  | 5-10 | 1.0846 | 0.55115 | 0.53483 | 391.6056 | -1.4061 | -1.3106
10-15 | 1.0704 | 0.57699 | 0.55129 | -0.12518 | 0.02189 | 0.012053 |  | 10-15 | 1.0515 | 0.5509 | 0.54346 | 391.7393 | -1.3515 | -1.3085
15-20 | 1.0767 | 0.5524 | 0.53856 | -0.086597 | 0.032606 | 0.0099832 |  | 15-20 | 1.0662 | 0.54268 | 0.53904 | 391.8189 | -1.2825 | -1.2906
20-25 | 1.0771 | 0.58459 | 0.54799 | -0.046282 | 0.1044 | 0.015962 |  | 20-25 | 1.0727 | 0.54763 | 0.54149 | 391.8433 | -1.5156 | -1.3231
25-30 | 1.0803 | 0.57805 | 0.54771 | -0.037703 | 0.1145 | 0.0089466 |  | 25-30 | 1.0838 | 0.58131 | 0.54457 | 391.8353 | -1.3305 | -1.3069
30-35 | 1.0797 | 0.58664 | 0.54554 | -0.026153 | -0.0048416 | -0.0012707 |  | 30-35 | 1.0828 | 0.56601 | 0.54376 | 391.8895 | -1.4654 | -1.3134
35-40 | 1.0801 | 0.55341 | 0.54025 | -0.064935 | 0.04094 | -0.019099 |  | 35-40 | 1.0679 | 0.5744 | 0.53805 | 391.8731 | -1.3588 | -1.2898
40-45 | 1.08 | 0.6031 | 0.53815 | -0.15361 | -0.0063662 | -0.017192 |  | 40-45 | 1.062 | 0.53497 | 0.54868 | 391.8712 | -1.4087 | -1.304
45-50 | 1.0734 | 0.59754 | 0.54435 | -0.19108 | -0.062039 | 0.009048 |  | 45-50 | 1.0847 | 0.56621 | 0.54739 | 391.9052 | -1.4396 | -1.2887
50-55 | 1.0804 | 0.54407 | 0.54262 | -0.16184 | -0.1063 | 0.0071234 |  | 50-55 | 1.0802 | 0.57311 | 0.548 | 391.8995 | -1.4012 | -1.3074
55-60 | 1.0737 | 0.57469 | 0.54715 | -0.20662 | 0.030497 | 0.015581 |  | 55-60 | 1.0785 | 0.56922 | 0.54213 | 391.9487 | -1.3624 | -1.2932
u(x) -- 0 N | [7 empty] | u(x) -- 40 kg
Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg) |  | Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg)
0-1 | - | - | - | 0.034095 | 0.01649 | 0.017055 |  | 0-1 | - | - | - | 0.033442 | 0.017453 | 0.017313
1-2 | - | - | - | 0.033722 | 0.016723 | 0.017111 |  | 1-2 | - | - | - | 0.034464 | 0.017112 | 0.016907
2-3 | - | - | - | 0.034164 | 0.019855 | 0.01747 |  | 2-3 | - | - | - | 0.033915 | 0.017971 | 0.017062
3-4 | - | - | - | 0.034191 | 0.016957 | 0.017387 |  | 3-4 | - | - | - | 0.034459 | 0.019052 | 0.017216
4-5 | - | - | - | 0.033628 | 0.018459 | 0.017127 |  | 4-5 | - | - | - | 0.033863 | 0.017076 | 0.017035
5-10 | - | - | - | 0.015088 | 0.0084662 | 0.0076663 |  | 5-10 | - | - | - | 0.015193 | 0.0079119 | 0.0076056
10-15 | - | - | - | 0.015118 | 0.0081242 | 0.007726 |  | 10-15 | - | - | - | 0.014882 | 0.0081536 | 0.0076351
15-20 | - | - | - | 0.015145 | 0.0080499 | 0.0076514 |  | 15-20 | - | - | - | 0.015193 | 0.0077864 | 0.0076742
20-25 | - | - | - | 0.015202 | 0.0084093 | 0.0076966 |  | 20-25 | - | - | - | 0.015192 | 0.0082842 | 0.0076578
25-30 | - | - | - | 0.015195 | 0.0083725 | 0.0077635 |  | 25-30 | - | - | - | 0.015152 | 0.0081134 | 0.0077173
30-35 | - | - | - | 0.015224 | 0.0084887 | 0.0077342 |  | 30-35 | - | - | - | 0.015144 | 0.0079452 | 0.0076156
35-40 | - | - | - | 0.015218 | 0.0081616 | 0.0076479 |  | 35-40 | - | - | - | 0.015113 | 0.0082428 | 0.007669
40-45 | - | - | - | 0.015184 | 0.0084905 | 0.0076513 |  | 40-45 | - | - | - | 0.014989 | 0.0079989 | 0.00771
45-50 | - | - | - | 0.015082 | 0.0081965 | 0.0076288 |  | 45-50 | - | - | - | 0.015192 | 0.0081369 | 0.0077238
50-55 | - | - | - | 0.015188 | 0.0083687 | 0.00765 |  | 50-55 | - | - | - | 0.015133 | 0.0079289 | 0.0076768
55-60 | - | - | - | 0.01544 | 0.0085654 | 0.0078728 |  | 55-60 | - | - | - | 0.015536 | 0.0082772 | 0.0078694

## REPETEABILITY TESTS
Pesas utilizadas |  | Código | Código | Código |  | Código | Código |  | 1 | Gravity (m/s^2)
pie izquierdo |  | M-015 | M-016 | M-019 |  | M-005 | M-006
pie derecho |  | M-022 | M-018 | M-020 |  | M-007
A 80 kg -- REP I | [7 empty] | A 140 kg -- REP I
Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg) |  | Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg)
0-1 | 1.0651 | 0.57216 | 0.55128 | 784.8046 | -2.3042 | -2.7536 |  | 0-1 | 1.0435 | 0.60183 | 0.5267 | 1373.2888 | -4.5255 | -4.774
1-2 | 1.0575 | 0.56668 | 0.5477 | 784.8661 | -2.6926 | -2.7089 |  | 1-2 | 0.99586 | 0.53455 | 0.54009 | 1373.3186 | -4.3474 | -4.8343
2-3 | 1.0671 | 0.55463 | 0.54734 | 784.8775 | -2.3237 | -2.7161 |  | 2-3 | 1.0486 | 0.54377 | 0.52462 | 1373.3595 | -4.2102 | -4.8061
3-4 | 1.0197 | 0.60334 | 0.53451 | 784.9833 | -2.4133 | -2.7363 |  | 3-4 | 1.0635 | 0.54685 | 0.51439 | 1373.4183 | -4.4108 | -4.8004
4-5 | 1.0335 | 0.64439 | 0.55344 | 785.054 | -2.3672 | -2.7213 |  | 4-5 | 1.0672 | 0.56067 | 0.52403 | 1373.3216 | -4.3966 | -4.8285
5-10 | 1.0434 | 0.5952 | 0.54423 | 785.2081 | -2.3555 | -2.7334 |  | 5-10 | 1.0515 | 0.53943 | 0.52825 | 1373.4825 | -4.3862 | -4.827
10-15 | 1.0601 | 0.58343 | 0.53536 | 785.4232 | -2.4303 | -2.7182 |  | 10-15 | 1.0433 | 0.5555 | 0.53118 | 1373.6565 | -4.3945 | -4.8388
15-20 | 1.0464 | 0.58195 | 0.53828 | 785.5269 | -2.3193 | -2.7286 |  | 15-20 | 1.0502 | 0.55959 | 0.53138 | 1373.9011 | -4.3702 | -4.8135
20-25 | 1.0306 | 0.5566 | 0.54711 | 785.5893 | -2.3068 | -2.7713 |  | 20-25 | 1.0462 | 0.55524 | 0.5332 | 1373.9204 | -4.2273 | -4.8305
25-30 | 1.0591 | 0.60061 | 0.53021 | 785.6907 | -2.4152 | -2.7805 |  | 25-30 | 1.0436 | 0.55176 | 0.52706 | 1374.1561 | -4.2953 | -4.8609
30-35 | 1.0319 | 0.61906 | 0.53247 | 785.8625 | -2.3572 | -2.7963 |  | 30-35 | 1.052 | 0.55215 | 0.52834 | 1374.0569 | -4.2885 | -4.8691
35-40 | 1.0423 | 0.56991 | 0.54309 | 785.9539 | -2.5021 | -2.8132 |  | 35-40 | 1.0342 | 0.52359 | 0.53089 | 1374.2734 | -4.3223 | -4.8683
40-45 | 1.0519 | 0.57471 | 0.5343 | 785.9831 | -2.3879 | -2.8341 |  | 40-45 | 1.0481 | 0.56301 | 0.5284 | 1374.3427 | -4.2811 | -4.8773
45-50 | 1.0408 | 0.57255 | 0.54036 | 786.0381 | -2.3418 | -2.8287 |  | 45-50 | 1.0463 | 0.54706 | 0.52958 | 1374.351 | -4.2261 | -4.9066
50-55 | 1.0312 | 0.58372 | 0.52423 | 786.113 | -2.3617 | -2.8277 |  | 50-55 | 1.0373 | 0.54011 | 0.52944 | 1374.4241 | -4.2936 | -4.895
55-60 | 1.0322 | 0.57863 | 0.54553 | 786.1822 | -2.3845 | -2.8087 |  | 55-60 | 1.0382 | 0.54972 | 0.53221 | 1374.489 | -4.2775 | -4.8829
A 80 kg -- REP II | [7 empty] | A 140 kg -- REP II
Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg) |  | Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg)
0-1 | 1.1482 | 0.57889 | 0.55082 | 783.8927 | -2.613 | -2.664 |  | 0-1 | 1.1135 | 0.51038 | 0.55642 | 1373.5625 | -4.2581 | -4.862
1-2 | 1.0717 | 0.57143 | 0.54971 | 783.8923 | -2.5029 | -2.7168 |  | 1-2 | 1.0866 | 0.5175 | 0.53983 | 1373.5912 | -4.2729 | -4.8308
2-3 | 1.0392 | 0.58345 | 0.54347 | 783.9975 | -2.5472 | -2.6787 |  | 2-3 | 1.0755 | 0.5406 | 0.54839 | 1373.6706 | -4.3886 | -4.8267
3-4 | 1.0441 | 0.60587 | 0.53854 | 784.0997 | -2.71 | -2.6616 |  | 3-4 | 1.0471 | 0.53301 | 0.54535 | 1373.7458 | -4.2828 | -4.8547
4-5 | 1.0513 | 0.63489 | 0.53828 | 784.1692 | -2.5137 | -2.6625 |  | 4-5 | 1.0668 | 0.504 | 0.53705 | 1373.8268 | -4.284 | -4.8352
5-10 | 1.0395 | 0.61026 | 0.54302 | 784.3519 | -2.5 | -2.6664 |  | 5-10 | 1.0702 | 0.54494 | 0.5439 | 1373.9145 | -4.2485 | -4.8507
10-15 | 1.0519 | 0.69504 | 0.54549 | 784.5334 | -2.5047 | -2.6832 |  | 10-15 | 1.0705 | 0.5572 | 0.54306 | 1374.0845 | -4.3481 | -4.8332
15-20 | 1.0485 | 0.58513 | 0.53987 | 784.6747 | -2.458 | -2.7014 |  | 15-20 | 1.072 | 0.53927 | 0.56953 | 1374.2063 | -4.3891 | -4.8704
20-25 | 1.052 | 0.60267 | 0.53776 | 784.8297 | -2.4178 | -2.6966 |  | 20-25 | 1.0761 | 0.5671 | 0.53952 | 1374.3581 | -4.4149 | -4.8735
25-30 | 1.06 | 0.58702 | 0.53834 | 784.9233 | -2.4781 | -2.6842 |  | 25-30 | 1.0775 | 0.56089 | 0.54868 | 1374.3945 | -4.3949 | -4.8868
30-35 | 1.0552 | 0.61199 | 0.54355 | 784.9528 | -2.4509 | -2.6694 |  | 30-35 | 1.1007 | 0.54923 | 0.54836 | 1374.4565 | -4.2971 | -4.8796
35-40 | 1.06 | 0.62863 | 0.53676 | 785.0717 | -2.3813 | -2.6547 |  | 35-40 | 1.0676 | 0.56127 | 0.53898 | 1374.5017 | -4.2648 | -4.8994
40-45 | 1.0436 | 0.60561 | 0.5388 | 785.0231 | -2.4911 | -2.6897 |  | 40-45 | 1.0587 | 0.52867 | 0.54525 | 1374.6311 | -4.1989 | -4.9082
45-50 | 1.0578 | 0.59031 | 0.53322 | 785.0837 | -2.6956 | -2.6832 |  | 45-50 | 1.069 | 0.54576 | 0.54988 | 1374.6123 | -4.2703 | -4.9
50-55 | 1.0549 | 0.60017 | 0.54409 | 785.1516 | -2.4585 | -2.6789 |  | 50-55 | 1.0509 | 0.54392 | 0.53881 | 1374.6166 | -4.2629 | -4.893
55-60 | 1.0526 | 0.58322 | 0.54129 | 785.0925 | -2.477 | -2.6774 |  | 55-60 | 1.0668 | 0.5564 | 0.54456 | 1374.6574 | -4.3161 | -4.899
A 80 kg -- REP III | [7 empty] | A 140 kg -- REP III
Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg) |  | Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg)
0-1 | 1.0616 | 0.55389 | 0.57076 | 784.4259 | -2.7845 | -2.7045 |  | 0-1 | 1.0763 | 0.52429 | 0.53256 | 1372.9919 | -4.3904 | -4.7969
1-2 | 1.0328 | 0.62992 | 0.54841 | 784.5143 | -2.6945 | -2.7251 |  | 1-2 | 1.1226 | 0.54748 | 0.55561 | 1373.0101 | -4.4227 | -4.8086
2-3 | 1.0614 | 0.59985 | 0.5475 | 784.5712 | -2.3472 | -2.7091 |  | 2-3 | 1.0037 | 0.5698 | 0.54707 | 1373.0433 | -4.3699 | -4.8337
3-4 | 1.0624 | 0.57636 | 0.54113 | 784.6185 | -2.5442 | -2.7451 |  | 3-4 | 1.1016 | 0.51964 | 0.56232 | 1373.1103 | -4.6016 | -4.8204
4-5 | 1.0755 | 0.5522 | 0.54719 | 784.7434 | -2.8624 | -2.7495 |  | 4-5 | 1.1 | 0.52063 | 0.53682 | 1373.2073 | -4.3866 | -4.826
5-10 | 1.0686 | 0.56023 | 0.54656 | 784.8582 | -2.7929 | -2.7331 |  | 5-10 | 1.0953 | 0.54767 | 0.54155 | 1373.1942 | -4.3969 | -4.8461
10-15 | 1.0619 | 0.5891 | 0.53575 | 785.0956 | -2.8561 | -2.7357 |  | 10-15 | 1.0834 | 0.56162 | 0.54855 | 1373.4332 | -4.3854 | -4.8343
15-20 | 1.0449 | 0.57799 | 0.54239 | 785.2349 | -2.8385 | -2.7353 |  | 15-20 | 1.1041 | 0.55599 | 0.5398 | 1373.5692 | -4.3827 | -4.8506
20-25 | 1.067 | 0.58285 | 0.5424 | 785.3548 | -2.6489 | -2.7474 |  | 20-25 | 1.111 | 0.5569 | 0.54216 | 1373.7236 | -4.4111 | -4.8545
25-30 | 1.057 | 0.61236 | 0.54618 | 785.4278 | -2.7809 | -2.7631 |  | 25-30 | 1.0739 | 0.54273 | 0.5437 | 1373.8435 | -4.4391 | -4.8621
30-35 | 1.0514 | 0.57827 | 0.54667 | 785.4497 | -2.7694 | -2.7701 |  | 30-35 | 1.088 | 0.57577 | 0.54671 | 1373.9418 | -4.4445 | -4.8773
35-40 | 1.0642 | 0.58064 | 0.53407 | 785.6032 | -2.6968 | -2.7491 |  | 35-40 | 1.0728 | 0.54329 | 0.55594 | 1374.0185 | -4.3522 | -4.8517
40-45 | 1.0515 | 0.55926 | 0.54593 | 785.5536 | -2.6885 | -2.7654 |  | 40-45 | 1.105 | 0.53397 | 0.54457 | 1374.0599 | -4.3567 | -4.8717
45-50 | 1.0474 | 0.57326 | 0.5449 | 785.7037 | -2.6598 | -2.7482 |  | 45-50 | 1.0867 | 0.55799 | 0.54497 | 1374.0634 | -4.3583 | -4.8766
50-55 | 1.0579 | 0.59935 | 0.54118 | 785.8552 | -2.7797 | -2.755 |  | 50-55 | 1.1121 | 0.5569 | 0.53459 | 1374.0899 | -4.3104 | -4.9045
55-60 | 1.0562 | 0.59085 | 0.53295 | 785.8428 | -2.7801 | -2.7821 |  | 55-60 | 1.099 | 0.5387 | 0.54056 | 1374.189 | -4.3817 | -4.901
A 80 kg -- REP IV | [7 empty] | A 140 kg -- REP IV
Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg) |  | Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg)
0-1 | 1.0621 | 0.56915 | 0.54384 | 784.1211 | -2.9802 | -2.6805 |  | 0-1 | 1.0354 | 0.50369 | 0.56315 | 1373.0291 | -4.3763 | -4.8022
1-2 | 1.0402 | 0.55139 | 0.54749 | 784.3069 | -2.5221 | -2.6593 |  | 1-2 | 1.0461 | 0.51508 | 0.53456 | 1373.1296 | -4.311 | -4.7632
2-3 | 1.0849 | 0.62675 | 0.5374 | 784.2951 | -2.4789 | -2.6879 |  | 2-3 | 1.0763 | 0.53407 | 0.5712 | 1373.1848 | -4.2084 | -4.7541
3-4 | 1.0837 | 0.54358 | 0.5477 | 784.347 | -2.7767 | -2.721 |  | 3-4 | 1.0593 | 0.51914 | 0.54388 | 1373.2621 | -4.5077 | -4.7822
4-5 | 1.071 | 0.55707 | 0.55724 | 784.3679 | -2.555 | -2.6873 |  | 4-5 | 1.0747 | 0.51492 | 0.53867 | 1373.3875 | -4.3467 | -4.7958
5-10 | 1.0543 | 0.5983 | 0.53856 | 784.5262 | -2.5742 | -2.7051 |  | 5-10 | 1.0709 | 0.52985 | 0.55639 | 1373.5127 | -4.2898 | -4.802
10-15 | 1.035 | 0.59659 | 0.54646 | 784.7145 | -2.5556 | -2.6865 |  | 10-15 | 1.0702 | 0.52747 | 0.55183 | 1373.7564 | -4.3804 | -4.8434
15-20 | 1.0629 | 0.60096 | 0.54139 | 784.9164 | -2.7281 | -2.7479 |  | 15-20 | 1.0609 | 0.53007 | 0.53798 | 1373.9494 | -4.3767 | -4.8328
20-25 | 1.0631 | 0.56179 | 0.54866 | 785.0363 | -2.584 | -2.7593 |  | 20-25 | 1.0694 | 0.52379 | 0.54103 | 1373.8298 | -4.3315 | -4.8487
25-30 | 1.0623 | 0.59899 | 0.53735 | 785.1283 | -2.7548 | -2.7817 |  | 25-30 | 1.0777 | 0.52734 | 0.54788 | 1374.0059 | -4.2513 | -4.8353
30-35 | 1.0531 | 0.58258 | 0.53992 | 785.1994 | -2.5977 | -2.764 |  | 30-35 | 1.0807 | 0.54429 | 0.54409 | 1374.201 | -4.2984 | -4.8696
35-40 | 1.0611 | 0.56013 | 0.54672 | 785.2662 | -2.7222 | -2.7806 |  | 35-40 | 1.0951 | 0.53306 | 0.5405 | 1374.288 | -4.3274 | -4.8596
40-45 | 1.0454 | 0.56703 | 0.55194 | 785.3265 | -2.7202 | -2.7889 |  | 40-45 | 1.0895 | 0.52941 | 0.54057 | 1374.3222 | -4.3883 | -4.8874
45-50 | 1.0575 | 0.59146 | 0.53826 | 785.4126 | -2.6455 | -2.7789 |  | 45-50 | 1.0909 | 0.53232 | 0.53973 | 1374.3952 | -4.311 | -4.9341
50-55 | 1.0695 | 0.59378 | 0.54187 | 785.4437 | -2.7878 | -2.8043 |  | 50-55 | 1.0916 | 0.53351 | 0.53862 | 1374.4824 | -4.2241 | -4.9349
55-60 | 1.057 | 0.58783 | 0.54153 | 785.4472 | -2.697 | -2.8144 |  | 55-60 | 1.0833 | 0.54075 | 0.54882 | 1374.521 | -4.158 | -4.929
A 80 kg -- REP V | [7 empty] | A 140 kg -- REP V
Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg) |  | Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg)
0-1 | 1.1378 | 0.57014 | 0.55671 | 784.4919 | -2.7946 | -2.7187 |  | 0-1 | 1.0891 | 0.53362 | 0.54784 | 1373.6099 | -4.4616 | -4.8483
1-2 | 1.0494 | 0.58127 | 0.53397 | 784.5223 | -2.8489 | -2.7026 |  | 1-2 | 1.1246 | 0.51097 | 0.54819 | 1373.6266 | -4.5698 | -4.8196
2-3 | 1.0748 | 0.54231 | 0.54915 | 784.5015 | -2.6314 | -2.737 |  | 2-3 | 1.1289 | 0.53842 | 0.57334 | 1373.7819 | -4.3225 | -4.8394
3-4 | 1.0584 | 0.57951 | 0.5797 | 784.5783 | -2.7485 | -2.7396 |  | 3-4 | 1.1001 | 0.50124 | 0.5394 | 1373.9356 | -4.4142 | -4.8095
4-5 | 1.0784 | 0.54885 | 0.53077 | 784.6959 | -3.0693 | -2.7278 |  | 4-5 | 1.081 | 0.52571 | 0.54792 | 1373.9164 | -4.4093 | -4.8394
5-10 | 1.0692 | 0.65577 | 0.5489 | 784.83 | -2.9026 | -2.7194 |  | 5-10 | 1.1014 | 0.54682 | 0.54604 | 1374.0855 | -4.4062 | -4.8844
10-15 | 1.0773 | 0.5662 | 0.53924 | 784.9819 | -2.932 | -2.7294 |  | 10-15 | 1.0806 | 0.54804 | 0.55468 | 1374.2366 | -4.4127 | -4.9096
15-20 | 1.0475 | 0.59706 | 0.53948 | 785.0196 | -2.8003 | -2.7082 |  | 15-20 | 1.0921 | 0.52909 | 0.55179 | 1374.2966 | -4.5003 | -4.9251
20-25 | 1.0402 | 0.58797 | 0.54627 | 785.2253 | -2.8967 | -2.7372 |  | 20-25 | 1.0901 | 0.52907 | 0.55236 | 1374.4855 | -4.3423 | -4.9002
25-30 | 1.0544 | 0.5747 | 0.53752 | 785.2861 | -2.8528 | -2.7542 |  | 25-30 | 1.0922 | 0.5434 | 0.54643 | 1374.5739 | -4.3187 | -4.8958
30-35 | 1.0689 | 0.60205 | 0.53666 | 785.4015 | -2.702 | -2.7786 |  | 30-35 | 1.0825 | 0.54178 | 0.54483 | 1374.5643 | -4.4225 | -4.9137
35-40 | 1.053 | 0.57259 | 0.53814 | 785.533 | -2.903 | -2.777 |  | 35-40 | 1.1002 | 0.53399 | 0.55627 | 1374.631 | -4.4532 | -4.9359
40-45 | 1.0616 | 0.57149 | 0.53688 | 785.5837 | -2.8892 | -2.7809 |  | 40-45 | 1.0675 | 0.53252 | 0.54386 | 1374.7863 | -4.4906 | -4.9624
45-50 | 1.0436 | 0.56049 | 0.54327 | 785.5737 | -2.8962 | -2.7844 |  | 45-50 | 1.1061 | 0.5473 | 0.54641 | 1374.7951 | -4.4859 | -4.9393
50-55 | 1.0788 | 0.58777 | 0.54548 | 785.5448 | -2.7744 | -2.7639 |  | 50-55 | 1.09 | 0.53404 | 0.54596 | 1374.7438 | -4.4108 | -4.9284
55-60 | 1.0589 | 0.62535 | 0.54576 | 785.6444 | -2.6789 | -2.7511 |  | 55-60 | 1.0743 | 0.5272 | 0.55192 | 1374.87 | -4.3571 | -4.9666
A 80 kg -- REP VI | [7 empty] | A 140 kg -- REP VI
Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg) |  | Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg)
0-1 | 1.0414 | 0.52883 | 0.55038 | 784.5734 | -2.7332 | -2.6659 |  | 0-1 | 1.078 | 0.52328 | 0.55312 | 1373.7469 | -4.8367 | -4.9858
1-2 | 1.0754 | 0.53595 | 0.53856 | 784.6664 | -2.2741 | -2.682 |  | 1-2 | 1.0759 | 0.52272 | 0.55373 | 1373.7968 | -4.849 | -4.9833
2-3 | 1.0702 | 0.54574 | 0.54491 | 784.6843 | -2.7608 | -2.6801 |  | 2-3 | 1.1014 | 0.56213 | 0.55788 | 1373.7542 | -4.6685 | -4.9767
3-4 | 1.1062 | 0.58082 | 0.56498 | 784.6652 | -2.7815 | -2.6902 |  | 3-4 | 1.0969 | 0.54705 | 0.57619 | 1373.8629 | -4.8048 | -5.003
4-5 | 1.0462 | 0.55026 | 0.55367 | 784.803 | -2.8254 | -2.702 |  | 4-5 | 1.0981 | 0.52591 | 0.57326 | 1373.8625 | -4.7712 | -4.9965
5-10 | 1.0847 | 0.58058 | 0.54826 | 784.8221 | -2.7722 | -2.7117 |  | 5-10 | 1.0912 | 0.55791 | 0.54984 | 1374.0337 | -4.734 | -4.9862
10-15 | 1.08 | 0.60796 | 0.53874 | 784.9627 | -2.5841 | -2.703 |  | 10-15 | 1.093 | 0.56616 | 0.55083 | 1374.3019 | -4.5383 | -4.9834
15-20 | 1.073 | 0.57437 | 0.54545 | 785.2721 | -2.6339 | -2.6952 |  | 15-20 | 1.0846 | 0.55086 | 0.54378 | 1374.5191 | -4.6518 | -5.0188
20-25 | 1.0611 | 0.55102 | 0.54526 | 785.3621 | -2.5834 | -2.6975 |  | 20-25 | 1.0812 | 0.54761 | 0.54356 | 1374.698 | -4.5357 | -5.0478
25-30 | 1.0469 | 0.57791 | 0.5404 | 785.4688 | -2.4596 | -2.6753 |  | 25-30 | 1.09 | 0.54592 | 0.55176 | 1374.8595 | -4.6288 | -5.062
30-35 | 1.0511 | 0.55943 | 0.54328 | 785.5067 | -2.5999 | -2.7059 |  | 30-35 | 1.0881 | 0.5436 | 0.54644 | 1374.9776 | -4.6326 | -5.0962
35-40 | 1.047 | 0.58422 | 0.54917 | 785.6047 | -2.6771 | -2.7143 |  | 35-40 | 1.086 | 0.54524 | 0.54645 | 1374.9873 | -4.6147 | -5.1019
40-45 | 1.0591 | 0.60212 | 0.53572 | 785.7553 | -2.6243 | -2.7106 |  | 40-45 | 1.094 | 0.57558 | 0.54846 | 1375.0597 | -4.6893 | -5.1028
45-50 | 1.0642 | 0.54818 | 0.5482 | 785.8265 | -2.6018 | -2.7024 |  | 45-50 | 1.0778 | 0.54325 | 0.5413 | 1375.2286 | -4.6679 | -5.0991
50-55 | 1.0641 | 0.57222 | 0.55233 | 785.8897 | -2.4591 | -2.6873 |  | 50-55 | 1.0979 | 0.55367 | 0.55081 | 1375.2738 | -4.6763 | -5.1174
55-60 | 1.0722 | 0.57601 | 0.54668 | 785.9578 | -2.5826 | -2.6942 |  | 55-60 | 1.0845 | 0.59051 | 0.5542 | 1375.3204 | -4.4631 | -5.1123
A 80 kg -- REP VII | [7 empty] | A 140 kg -- REP VII
Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg) |  | Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg)
0-1 | 1.0775 | 0.5635 | 0.53251 | 784.5194 | -2.6401 | -2.6915 |  | 0-1 | 1.0615 | 0.52228 | 0.53983 | 1374.1702 | -5.241 | -4.968
1-2 | 1.0873 | 0.50665 | 0.55452 | 784.5795 | -2.6141 | -2.6743 |  | 1-2 | 1.0724 | 0.55268 | 0.54899 | 1374.2035 | -4.8905 | -4.9397
2-3 | 1.0832 | 0.53126 | 0.53102 | 784.6727 | -2.8886 | -2.6748 |  | 2-3 | 1.062 | 0.52482 | 0.54914 | 1374.2292 | -4.6059 | -4.956
3-4 | 1.0729 | 0.52713 | 0.5509 | 784.6724 | -2.3922 | -2.7152 |  | 3-4 | 1.096 | 0.53456 | 0.54883 | 1374.2898 | -4.8763 | -4.9825
4-5 | 1.0367 | 0.53236 | 0.56964 | 784.7509 | -2.653 | -2.6906 |  | 4-5 | 1.0692 | 0.54678 | 0.55331 | 1374.3847 | -4.7367 | -4.9311
5-10 | 1.0812 | 0.60022 | 0.54365 | 784.7765 | -2.6594 | -2.7291 |  | 5-10 | 1.0869 | 0.56879 | 0.5501 | 1374.5229 | -4.8815 | -4.9566
10-15 | 1.0473 | 0.58621 | 0.54411 | 784.8563 | -2.6605 | -2.7667 |  | 10-15 | 1.0943 | 0.61267 | 0.54172 | 1374.7932 | -4.8231 | -4.9704
15-20 | 1.0719 | 0.55735 | 0.5433 | 785.0432 | -2.8186 | -2.7606 |  | 15-20 | 1.087 | 0.5724 | 0.54779 | 1374.8162 | -4.8917 | -4.9892
20-25 | 1.0717 | 0.56161 | 0.54585 | 785.1263 | -2.5745 | -2.7687 |  | 20-25 | 1.0696 | 0.55821 | 0.54162 | 1374.935 | -5.0895 | -5.005
25-30 | 1.0794 | 0.58975 | 0.54558 | 785.2591 | -2.7706 | -2.789 |  | 25-30 | 1.0833 | 0.58658 | 0.54834 | 1375.1366 | -4.7927 | -5.0019
30-35 | 1.0638 | 0.58268 | 0.55332 | 785.3125 | -2.8162 | -2.7868 |  | 30-35 | 1.0779 | 0.59271 | 0.54912 | 1375.2036 | -4.7819 | -5.011
35-40 | 1.0594 | 0.588 | 0.54853 | 785.3996 | -2.6566 | -2.794 |  | 35-40 | 1.0661 | 0.5789 | 0.54745 | 1375.4147 | -4.6434 | -4.9951
40-45 | 1.072 | 0.56203 | 0.54626 | 785.4763 | -2.7832 | -2.8114 |  | 40-45 | 1.0764 | 0.61132 | 0.55377 | 1375.3472 | -4.8716 | -5.0134
45-50 | 1.0506 | 0.60561 | 0.54649 | 785.585 | -2.8227 | -2.8166 |  | 45-50 | 1.0739 | 0.59108 | 0.55401 | 1375.5303 | -4.7834 | -5.0516
50-55 | 1.0666 | 0.60929 | 0.54281 | 785.675 | -2.7718 | -2.8483 |  | 50-55 | 1.0711 | 0.59622 | 0.55359 | 1375.5629 | -4.9103 | -5.0351
55-60 | 1.0605 | 0.62005 | 0.5387 | 785.6959 | -2.6523 | -2.8205 |  | 55-60 | 1.0769 | 0.60434 | 0.55689 | 1375.4925 | -4.913 | -5.0552
A 80 kg -- REP VIII | [7 empty] | A 140 kg -- REP VIII
Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg) |  | Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg)
0-1 | 1.0591 | 0.5747 | 0.54238 | 784.1608 | -3.0103 | -2.6363 |  | 0-1 | 1.0649 | 0.59548 | 0.53041 | 1373.8627 | -4.64 | -4.898
1-2 | 1.0762 | 0.55324 | 0.54651 | 784.1844 | -2.7409 | -2.6467 |  | 1-2 | 1.0692 | 0.62716 | 0.53742 | 1373.8919 | -4.3172 | -4.928
2-3 | 1.085 | 0.63129 | 0.55161 | 784.2738 | -2.7833 | -2.6452 |  | 2-3 | 1.0763 | 0.58832 | 0.53966 | 1373.9334 | -4.4431 | -4.9011
3-4 | 1.0636 | 0.56926 | 0.55918 | 784.2783 | -2.8016 | -2.6438 |  | 3-4 | 1.0903 | 0.61928 | 0.54966 | 1373.9627 | -4.6005 | -4.932
4-5 | 1.0791 | 0.58967 | 0.55233 | 784.4257 | -2.7396 | -2.6157 |  | 4-5 | 1.125 | 0.5306 | 0.54441 | 1374.0515 | -4.6946 | -4.908
5-10 | 1.0715 | 0.54652 | 0.54757 | 784.568 | -2.7981 | -2.6254 |  | 5-10 | 1.0576 | 0.63747 | 0.55168 | 1374.2431 | -4.4172 | -4.9243
10-15 | 1.0772 | 0.5627 | 0.54825 | 784.7193 | -2.8005 | -2.6422 |  | 10-15 | 1.0782 | 0.59779 | 0.55043 | 1374.327 | -4.617 | -4.9347
15-20 | 1.0553 | 0.59607 | 0.55252 | 784.8878 | -2.6832 | -2.6445 |  | 15-20 | 1.0876 | 0.68536 | 0.55087 | 1374.5562 | -4.4276 | -4.9431
20-25 | 1.0636 | 0.55186 | 0.55703 | 785.0289 | -2.6446 | -2.6425 |  | 20-25 | 1.0907 | 0.68769 | 0.54625 | 1374.631 | -4.6008 | -4.9491
25-30 | 1.0911 | 0.58615 | 0.55364 | 785.0123 | -2.7392 | -2.6517 |  | 25-30 | 1.0824 | 0.6878 | 0.55638 | 1374.688 | -4.5346 | -4.938
30-35 | 1.0812 | 0.57636 | 0.54218 | 784.9865 | -2.7826 | -2.6657 |  | 30-35 | 1.0748 | 0.73869 | 0.54701 | 1374.7542 | -4.243 | -4.9652
35-40 | 1.0574 | 0.57079 | 0.54739 | 785.1001 | -2.7922 | -2.6638 |  | 35-40 | 1.0747 | 0.64268 | 0.55037 | 1374.8546 | -4.4091 | -4.9668
40-45 | 1.0736 | 0.60105 | 0.55526 | 785.0846 | -2.7862 | -2.6799 |  | 40-45 | 1.0758 | 0.60332 | 0.54692 | 1374.9732 | -4.4412 | -4.9611
45-50 | 1.0671 | 0.59565 | 0.54707 | 785.3817 | -2.8278 | -2.6729 |  | 45-50 | 1.0883 | 0.62379 | 0.54038 | 1375.0267 | -4.4721 | -4.9786
50-55 | 1.0629 | 0.61081 | 0.54983 | 785.3684 | -2.8311 | -2.6723 |  | 50-55 | 1.0884 | 0.70484 | 0.54679 | 1375.0904 | -4.288 | -4.9811
55-60 | 1.0753 | 0.55464 | 0.54063 | 785.4281 | -2.7551 | -2.733 |  | 55-60 | 1.0929 | 0.75819 | 0.54518 | 1375.1632 | -4.2525 | -4.9955
A 80 kg -- REP IX | [7 empty] | A 140 kg -- REP IX
Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg) |  | Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg)
0-1 | 1.1101 | 0.54215 | 0.54858 | 783.9908 | -2.5166 | -2.6727 |  | 0-1 | 1.1116 | 0.8172 | 0.56222 | 1373.2378 | -4.9435 | -4.9083
1-2 | 1.0503 | 0.56179 | 0.55356 | 784.1846 | -2.7465 | -2.6666 |  | 1-2 | 1.1298 | 0.65406 | 0.55257 | 1373.3276 | -5.1384 | -4.9494
2-3 | 1.0705 | 0.54554 | 0.54462 | 784.2241 | -2.9787 | -2.6871 |  | 2-3 | 1.0856 | 0.52786 | 0.55485 | 1373.4123 | -5.1718 | -4.9381
3-4 | 1.0528 | 0.68506 | 0.54556 | 784.274 | -2.6691 | -2.678 |  | 3-4 | 1.1049 | 0.54847 | 0.56733 | 1373.5075 | -5.2288 | -4.9773
4-5 | 1.0616 | 0.50714 | 0.54272 | 784.3065 | -2.4915 | -2.6705 |  | 4-5 | 1.0879 | 0.60155 | 0.56106 | 1373.5294 | -5.4095 | -4.972
5-10 | 1.0703 | 0.54014 | 0.55188 | 784.4323 | -2.582 | -2.6756 |  | 5-10 | 1.0702 | 0.64536 | 0.5516 | 1373.7386 | -5.1327 | -4.9836
10-15 | 1.0801 | 0.57255 | 0.54656 | 784.6752 | -2.5316 | -2.6996 |  | 10-15 | 1.0678 | 0.66794 | 0.5478 | 1373.9059 | -5.1975 | -4.9639
15-20 | 1.0654 | 0.55902 | 0.54658 | 784.8643 | -2.8078 | -2.6897 |  | 15-20 | 1.104 | 0.64362 | 0.54646 | 1373.997 | -5.2236 | -4.9398
20-25 | 1.0771 | 0.58897 | 0.54012 | 784.865 | -2.6922 | -2.6903 |  | 20-25 | 1.1078 | 0.607 | 0.56277 | 1374.162 | -5.1231 | -4.983
25-30 | 1.0627 | 0.55094 | 0.55022 | 784.9099 | -2.541 | -2.7242 |  | 25-30 | 1.0922 | 0.66272 | 0.54482 | 1374.3263 | -5.2285 | -4.9871
30-35 | 1.07 | 0.54865 | 0.54881 | 785.0884 | -2.7004 | -2.7208 |  | 30-35 | 1.0884 | 0.65668 | 0.55696 | 1374.4976 | -5.5153 | -5.0143
35-40 | 1.0644 | 0.54901 | 0.54294 | 785.1903 | -2.5647 | -2.7536 |  | 35-40 | 1.1555 | 0.63645 | 0.5652 | 1374.5789 | -5.1563 | -5.0267
40-45 | 1.0708 | 0.55192 | 0.54809 | 785.269 | -2.627 | -2.7453 |  | 40-45 | 1.0781 | 0.63068 | 0.54691 | 1374.8253 | -5.2051 | -4.9946
45-50 | 1.0637 | 0.56062 | 0.55035 | 785.2862 | -2.6405 | -2.7304 |  | 45-50 | 1.0842 | 0.79033 | 0.54416 | 1374.554 | -5.0887 | -4.9931
50-55 | 1.0706 | 0.58045 | 0.55014 | 785.2227 | -2.5304 | -2.7276 |  | 50-55 | 1.0992 | 0.63771 | 0.54741 | 1374.7915 | -5.1573 | -4.9945
55-60 | 1.0756 | 0.57163 | 0.54016 | 785.312 | -2.6498 | -2.7544 |  | 55-60 | 1.0835 | 0.65636 | 0.54552 | 1374.7223 | -4.9532 | -4.9879
A 80 kg -- REP X | [7 empty] | A 140 kg -- REP X
Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg) |  | Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg)
0-1 | 1.0857 | 0.53842 | 0.54294 | 784.681 | -2.7556 | -2.7367 |  | 0-1 | 1.0485 | 0.52143 | 0.54737 | 1375.0768 | -4.8947 | -4.9365
1-2 | 1.0548 | 0.55565 | 0.55993 | 784.6753 | -2.5968 | -2.7369 |  | 1-2 | 1.0445 | 0.51605 | 0.53335 | 1374.9909 | -5.3787 | -4.9881
2-3 | 1.0713 | 0.53838 | 0.54663 | 784.6338 | -2.6666 | -2.7754 |  | 2-3 | 1.0695 | 0.6527 | 0.53919 | 1375.0693 | -5.0488 | -4.9539
3-4 | 1.0594 | 0.53796 | 0.56929 | 784.7434 | -2.5512 | -2.7899 |  | 3-4 | 1.0707 | 0.60772 | 0.53266 | 1375.0637 | -5.45 | -4.9923
4-5 | 1.0686 | 0.51963 | 0.5483 | 784.825 | -2.571 | -2.7873 |  | 4-5 | 1.093 | 0.52463 | 0.54925 | 1375.1427 | -5.0247 | -4.9508
5-10 | 1.0796 | 0.5689 | 0.54832 | 784.9801 | -2.534 | -2.7809 |  | 5-10 | 1.1038 | 0.58786 | 0.52349 | 1375.2554 | -5.214 | -4.9557
10-15 | 1.0636 | 0.56604 | 0.54608 | 785.2 | -2.5711 | -2.8071 |  | 10-15 | 1.0716 | 0.58624 | 0.54038 | 1375.4054 | -5.1355 | -4.959
15-20 | 1.0593 | 0.56398 | 0.54638 | 785.3356 | -2.5662 | -2.8238 |  | 15-20 | 1.0742 | 0.56435 | 0.53305 | 1375.6136 | -4.8677 | -4.9946
20-25 | 1.0585 | 0.5696 | 0.55003 | 785.3737 | -2.6303 | -2.8303 |  | 20-25 | 1.0644 | 0.5832 | 0.53958 | 1375.7769 | -4.9733 | -5.0055
25-30 | 1.0681 | 0.54793 | 0.54498 | 785.5691 | -2.5261 | -2.7911 |  | 25-30 | 1.0538 | 0.57882 | 0.53194 | 1375.8574 | -4.9679 | -4.9952
30-35 | 1.0821 | 0.54256 | 0.54748 | 785.596 | -2.614 | -2.7959 |  | 30-35 | 1.0535 | 0.61006 | 0.53591 | 1376.0441 | -5.0271 | -4.9929
35-40 | 1.0742 | 0.55564 | 0.5457 | 785.6902 | -2.6437 | -2.8056 |  | 35-40 | 1.0546 | 0.60876 | 0.53318 | 1376.1646 | -5.1748 | -5.01
40-45 | 1.0838 | 0.56181 | 0.54394 | 785.6644 | -2.5854 | -2.8191 |  | 40-45 | 1.0597 | 0.58302 | 0.52786 | 1376.2691 | -5.0179 | -5.0286
45-50 | 1.0765 | 0.60412 | 0.55014 | 785.7566 | -2.5269 | -2.8367 |  | 45-50 | 1.0632 | 0.55638 | 0.54429 | 1376.2532 | -4.887 | -5.0357
50-55 | 1.0737 | 0.58464 | 0.55048 | 785.8706 | -2.5096 | -2.8343 |  | 50-55 | 1.078 | 0.56084 | 0.52754 | 1376.3303 | -4.9441 | -5.0606
55-60 | 1.0611 | 0.56036 | 0.53954 | 785.9098 | -2.4835 | -2.8463 |  | 55-60 | 1.0608 | 0.6036 | 0.5343 | 1376.4981 | -4.8705 | -5.0809
A 80 kg -- REP MAX | [7 empty] | A 140 kg -- REP MAX
Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg) |  | Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg)
0-1 | 1.1482 | 0.57889 | 0.57076 | 784.8046 | -2.3042 | -2.6363 |  | 0-1 | 1.1135 | 0.8172 | 0.56315 | 1375.0768 | -4.2581 | -4.774
1-2 | 1.0873 | 0.62992 | 0.55993 | 784.8661 | -2.2741 | -2.6467 |  | 1-2 | 1.1298 | 0.65406 | 0.55561 | 1374.9909 | -4.2729 | -4.7632
2-3 | 1.085 | 0.63129 | 0.55161 | 784.8775 | -2.3237 | -2.6452 |  | 2-3 | 1.1289 | 0.6527 | 0.57334 | 1375.0693 | -4.2084 | -4.7541
3-4 | 1.1062 | 0.68506 | 0.5797 | 784.9833 | -2.3922 | -2.6438 |  | 3-4 | 1.1049 | 0.61928 | 0.57619 | 1375.0637 | -4.2828 | -4.7822
4-5 | 1.0791 | 0.64439 | 0.56964 | 785.054 | -2.3672 | -2.6157 |  | 4-5 | 1.125 | 0.60155 | 0.57326 | 1375.1427 | -4.284 | -4.7958
5-10 | 1.0847 | 0.65577 | 0.55188 | 785.2081 | -2.3555 | -2.6254 |  | 5-10 | 1.1038 | 0.64536 | 0.55639 | 1375.2554 | -4.2485 | -4.802
10-15 | 1.0801 | 0.69504 | 0.54825 | 785.4232 | -2.4303 | -2.6422 |  | 10-15 | 1.0943 | 0.66794 | 0.55468 | 1375.4054 | -4.3481 | -4.8332
15-20 | 1.073 | 0.60096 | 0.55252 | 785.5269 | -2.3193 | -2.6445 |  | 15-20 | 1.1041 | 0.68536 | 0.56953 | 1375.6136 | -4.3702 | -4.8135
20-25 | 1.0771 | 0.60267 | 0.55703 | 785.5893 | -2.3068 | -2.6425 |  | 20-25 | 1.111 | 0.68769 | 0.56277 | 1375.7769 | -4.2273 | -4.8305
25-30 | 1.0911 | 0.61236 | 0.55364 | 785.6907 | -2.4152 | -2.6517 |  | 25-30 | 1.0922 | 0.6878 | 0.55638 | 1375.8574 | -4.2513 | -4.8353
30-35 | 1.0821 | 0.61906 | 0.55332 | 785.8625 | -2.3572 | -2.6657 |  | 30-35 | 1.1007 | 0.73869 | 0.55696 | 1376.0441 | -4.243 | -4.8691
35-40 | 1.0742 | 0.62863 | 0.54917 | 785.9539 | -2.3813 | -2.6547 |  | 35-40 | 1.1555 | 0.64268 | 0.5652 | 1376.1646 | -4.2648 | -4.8517
40-45 | 1.0838 | 0.60561 | 0.55526 | 785.9831 | -2.3879 | -2.6799 |  | 40-45 | 1.105 | 0.63068 | 0.55377 | 1376.2691 | -4.1989 | -4.8717
45-50 | 1.0765 | 0.60561 | 0.55035 | 786.0381 | -2.3418 | -2.6729 |  | 45-50 | 1.1061 | 0.79033 | 0.55401 | 1376.2532 | -4.2261 | -4.8766
50-55 | 1.0788 | 0.61081 | 0.55233 | 786.113 | -2.3617 | -2.6723 |  | 50-55 | 1.1121 | 0.70484 | 0.55359 | 1376.3303 | -4.2241 | -4.893
55-60 | 1.0756 | 0.62535 | 0.54668 | 786.1822 | -2.3845 | -2.6774 |  | 55-60 | 1.099 | 0.75819 | 0.55689 | 1376.4981 | -4.158 | -4.8829
A 80 kg -- REP MIN | [7 empty] | A 140 kg -- REP MIN
Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg) |  | Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg)
0-1 | 1.0414 | 0.52883 | 0.53251 | 783.8927 | -3.0103 | -2.7367 |  | 0-1 | 1.0354 | 0.50369 | 0.53041 | 1372.9919 | -5.241 | -4.9858
1-2 | 1.0328 | 0.50665 | 0.53397 | 783.8923 | -2.8489 | -2.7369 |  | 1-2 | 1.0445 | 0.51097 | 0.53335 | 1373.0101 | -5.3787 | -4.9881
2-3 | 1.0392 | 0.53126 | 0.53102 | 783.9975 | -2.9787 | -2.7754 |  | 2-3 | 1.0037 | 0.52482 | 0.53919 | 1373.0433 | -5.1718 | -4.9767
3-4 | 1.0441 | 0.52713 | 0.53854 | 784.0997 | -2.8016 | -2.7899 |  | 3-4 | 1.0471 | 0.50124 | 0.53266 | 1373.1103 | -5.45 | -5.003
4-5 | 1.0367 | 0.50714 | 0.53077 | 784.1692 | -3.0693 | -2.7873 |  | 4-5 | 1.0668 | 0.504 | 0.53682 | 1373.2073 | -5.4095 | -4.9965
5-10 | 1.0395 | 0.54014 | 0.53856 | 784.3519 | -2.9026 | -2.7809 |  | 5-10 | 1.0576 | 0.52985 | 0.52349 | 1373.1942 | -5.214 | -4.9862
10-15 | 1.035 | 0.5627 | 0.53575 | 784.5334 | -2.932 | -2.8071 |  | 10-15 | 1.0678 | 0.52747 | 0.54038 | 1373.4332 | -5.1975 | -4.9834
15-20 | 1.0449 | 0.55735 | 0.53948 | 784.6747 | -2.8385 | -2.8238 |  | 15-20 | 1.0609 | 0.52909 | 0.53305 | 1373.5692 | -5.2236 | -5.0188
20-25 | 1.0402 | 0.55102 | 0.53776 | 784.8297 | -2.8967 | -2.8303 |  | 20-25 | 1.0644 | 0.52379 | 0.53952 | 1373.7236 | -5.1231 | -5.0478
25-30 | 1.0469 | 0.54793 | 0.53735 | 784.9099 | -2.8528 | -2.7911 |  | 25-30 | 1.0538 | 0.52734 | 0.53194 | 1373.8435 | -5.2285 | -5.062
30-35 | 1.0511 | 0.54256 | 0.53666 | 784.9528 | -2.8162 | -2.7959 |  | 30-35 | 1.0535 | 0.54178 | 0.53591 | 1373.9418 | -5.5153 | -5.0962
35-40 | 1.047 | 0.54901 | 0.53407 | 785.0717 | -2.903 | -2.8056 |  | 35-40 | 1.0546 | 0.53306 | 0.53318 | 1374.0185 | -5.1748 | -5.1019
40-45 | 1.0436 | 0.55192 | 0.53572 | 785.0231 | -2.8892 | -2.8191 |  | 40-45 | 1.0587 | 0.52867 | 0.52786 | 1374.0599 | -5.2051 | -5.1028
45-50 | 1.0436 | 0.54818 | 0.53322 | 785.0837 | -2.8962 | -2.8367 |  | 45-50 | 1.0632 | 0.53232 | 0.53973 | 1374.0634 | -5.0887 | -5.0991
50-55 | 1.0549 | 0.57222 | 0.54118 | 785.1516 | -2.8311 | -2.8483 |  | 50-55 | 1.0509 | 0.53351 | 0.52754 | 1374.0899 | -5.1573 | -5.1174
55-60 | 1.0526 | 0.55464 | 0.53295 | 785.0925 | -2.7801 | -2.8463 |  | 55-60 | 1.0608 | 0.5272 | 0.5343 | 1374.189 | -4.9532 | -5.1123
A 80 kg -- REP average | [7 empty] | A 140 kg -- REP average
Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg) |  | Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg)
0-1 | 1.0825 | 0.55485 | 0.54914 | 784.3662 | -2.7529 | -2.688 |  | 0-1 | 1.0714 | 0.58622 | 0.54701 | 1373.7794 | -4.7283 | -4.8904
1-2 | 1.0587 | 0.56124 | 0.54769 | 784.4392 | -2.6161 | -2.6877 |  | 1-2 | 1.0859 | 0.56112 | 0.54534 | 1373.7978 | -4.7529 | -4.8931
2-3 | 1.0726 | 0.57237 | 0.54355 | 784.4731 | -2.6838 | -2.7017 |  | 2-3 | 1.0736 | 0.56756 | 0.55449 | 1373.8521 | -4.6219 | -4.8884
3-4 | 1.071 | 0.58119 | 0.55767 | 784.526 | -2.6459 | -2.7157 |  | 3-4 | 1.0872 | 0.55176 | 0.55291 | 1373.9169 | -4.8217 | -4.9084
4-5 | 1.0633 | 0.55087 | 0.55023 | 784.6141 | -2.7204 | -2.7034 |  | 4-5 | 1.0921 | 0.53963 | 0.55148 | 1373.9832 | -4.7473 | -4.9012
5-10 | 1.0704 | 0.58466 | 0.54641 | 784.7353 | -2.6873 | -2.7087 |  | 5-10 | 1.0839 | 0.57969 | 0.54506 | 1374.1036 | -4.6935 | -4.9127
10-15 | 1.0638 | 0.59051 | 0.54292 | 784.9162 | -2.6854 | -2.722 |  | 10-15 | 1.0801 | 0.58633 | 0.54813 | 1374.2998 | -4.7035 | -4.9215
15-20 | 1.0598 | 0.57851 | 0.54495 | 785.0776 | -2.7034 | -2.7273 |  | 15-20 | 1.086 | 0.58462 | 0.54541 | 1374.45 | -4.6916 | -4.9326
20-25 | 1.0619 | 0.57093 | 0.54704 | 785.1791 | -2.6458 | -2.7346 |  | 20-25 | 1.0859 | 0.58049 | 0.54716 | 1374.5742 | -4.6758 | -4.9472
25-30 | 1.066 | 0.5799 | 0.54469 | 785.2662 | -2.6693 | -2.7373 |  | 25-30 | 1.0792 | 0.58905 | 0.54596 | 1374.6992 | -4.6641 | -4.9475
30-35 | 1.0655 | 0.57342 | 0.54483 | 785.3356 | -2.6755 | -2.7449 |  | 30-35 | 1.0788 | 0.60841 | 0.54639 | 1374.817 | -4.7123 | -4.9705
35-40 | 1.0602 | 0.57387 | 0.54359 | 785.4413 | -2.6941 | -2.7498 |  | 35-40 | 1.0915 | 0.57981 | 0.54938 | 1374.9121 | -4.6571 | -4.9701
40-45 | 1.0645 | 0.57343 | 0.5455 | 785.472 | -2.6981 | -2.7601 |  | 40-45 | 1.081 | 0.57592 | 0.54346 | 1374.9972 | -4.6865 | -4.9797
45-50 | 1.0591 | 0.57932 | 0.54523 | 785.5648 | -2.6859 | -2.758 |  | 45-50 | 1.084 | 0.60651 | 0.5449 | 1375.0163 | -4.6369 | -4.9884
50-55 | 1.0678 | 0.59214 | 0.54676 | 785.6135 | -2.6637 | -2.7614 |  | 50-55 | 1.0891 | 0.59161 | 0.54264 | 1375.0785 | -4.6303 | -4.9967
55-60 | 1.0645 | 0.58667 | 0.54056 | 785.6513 | -2.6444 | -2.772 |  | 55-60 | 1.0815 | 0.6105 | 0.54686 | 1375.1464 | -4.546 | -5.0024
stdev -- 80 kg | [7 empty] | stdev -- 140 kg | [6 empty] | CRITICAL |  |  | u(x)-- repeatibility
Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg) |  | Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg) | AVERAGE FZ(kg) | AVERAGE FY (kg) | AVERAGE FX(kg) | AVERAGE FZ(kg) | AVERAGE FY (kg) | AVERAGE FX(kg)
0-1 | - | - | - | 0.30631 | 0.20944 | 0.032942 |  | 0-1 | - | - | - | 0.74885 | 0.33519 | 0.073581 | 0.74885 | 0.33519 | 0.073581 | 0.52952 | 0.23702 | 0.05203
1-2 | - | - | - | 0.29216 | 0.19347 | 0.032781 |  | 1-2 | - | - | - | 0.69497 | 0.39926 | 0.089089 | 0.69497 | 0.39926 | 0.089089 | 0.49142 | 0.28232 | 0.062995
2-3 | - | - | - | 0.26745 | 0.21928 | 0.039886 |  | 2-3 | - | - | - | 0.70022 | 0.33454 | 0.083206 | 0.70022 | 0.33454 | 0.083206 | 0.49513 | 0.23655 | 0.058836
3-4 | - | - | - | 0.26816 | 0.15708 | 0.047896 |  | 3-4 | - | - | - | 0.66756 | 0.39019 | 0.092393 | 0.66756 | 0.39019 | 0.092393 | 0.47204 | 0.27591 | 0.065332
4-5 | - | - | - | 0.27979 | 0.20989 | 0.054538 |  | 4-5 | - | - | - | 0.66856 | 0.36913 | 0.075216 | 0.66856 | 0.36913 | 0.075216 | 0.47274 | 0.26101 | 0.053185
5-10 | - | - | - | 0.26401 | 0.16804 | 0.049221 |  | 5-10 | - | - | - | 0.6788 | 0.36717 | 0.071548 | 0.6788 | 0.36717 | 0.071548 | 0.47998 | 0.25963 | 0.050592
10-15 | - | - | - | 0.2716 | 0.1657 | 0.051957 |  | 10-15 | - | - | - | 0.65813 | 0.32497 | 0.063935 | 0.65813 | 0.32497 | 0.063935 | 0.46537 | 0.22979 | 0.045209
15-20 | - | - | - | 0.25935 | 0.17315 | 0.054375 |  | 15-20 | - | - | - | 0.67758 | 0.29471 | 0.072463 | 0.67758 | 0.29471 | 0.072463 | 0.47912 | 0.20839 | 0.051239
20-25 | - | - | - | 0.24474 | 0.15711 | 0.058746 |  | 20-25 | - | - | - | 0.70739 | 0.33397 | 0.077933 | 0.70739 | 0.33397 | 0.077933 | 0.5002 | 0.23616 | 0.055107
25-30 | - | - | - | 0.27234 | 0.16213 | 0.056608 |  | 25-30 | - | - | - | 0.69365 | 0.32647 | 0.078712 | 0.69365 | 0.32647 | 0.078712 | 0.49048 | 0.23085 | 0.055658
30-35 | - | - | - | 0.28688 | 0.14735 | 0.052962 |  | 30-35 | - | - | - | 0.71534 | 0.41455 | 0.078762 | 0.71534 | 0.41455 | 0.078762 | 0.50582 | 0.29313 | 0.055693
35-40 | - | - | - | 0.28565 | 0.16446 | 0.058941 |  | 35-40 | - | - | - | 0.73954 | 0.34137 | 0.085061 | 0.73954 | 0.34137 | 0.085061 | 0.52293 | 0.24139 | 0.060147
40-45 | - | - | - | 0.30029 | 0.14851 | 0.05394 |  | 40-45 | - | - | - | 0.73578 | 0.34695 | 0.076515 | 0.73578 | 0.34695 | 0.076515 | 0.52027 | 0.24533 | 0.054104
45-50 | - | - | - | 0.28183 | 0.16135 | 0.060024 |  | 45-50 | - | - | - | 0.7497 | 0.29041 | 0.076604 | 0.7497 | 0.29041 | 0.076604 | 0.53012 | 0.20535 | 0.054167
50-55 | - | - | - | 0.31863 | 0.17752 | 0.067638 |  | 50-55 | - | - | - | 0.7542 | 0.3511 | 0.07734 | 0.7542 | 0.3511 | 0.07734 | 0.5333 | 0.24826 | 0.054688
55-60 | - | - | - | 0.33336 | 0.12943 | 0.060657 |  | 55-60 | - | - | - | 0.7878 | 0.31177 | 0.080717 | 0.7878 | 0.31177 | 0.080717 | 0.55706 | 0.22045 | 0.057075

## LINEARITY TESTS
UTILIZED WEIGHTS |  | CODE | CODE | CODE | CODE | CODE |  |  | u(x) digital resolution |  | 2.8868e-07
LEFT FOOT |  | M-015 | M-016 | M-019 | M-005 | M-006
NUMBER
[9 empty] | 1 | Gravity (m/s^2)
RIGHT FOOT |  | M-022 | M-018 | M-020 | M-007 | [28 empty] | 1
NUMBER
Gravedad at lab
LOAD | [8 empty] | gl | 9.7781 | [5 empty] | UNLOAD | [6 empty] | Io (INDICATION WITHOUT LOAD) | [6 empty] | Combined uncertainty for indication | [8 empty] | Incert. Exp indic |  |  | INDICATION ERROR
OIML127
0 kg -- REP UNIQUE | [8 empty] | Latitude (◦) | 0.17356 | [5 empty] | 0 kg -- REP UNIQUE | [6 empty] | 0 kg -- REP UNIQUE | [6 empty] | 0 kg -- REP UNIQUE | [8 empty] | A 0 kg -- REP unica
Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg) |  |  | Altitude (m) | 1225.9812 | [5 empty] | Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg) | Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg) | Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg) |  |  | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg)
0-1 | 1.0452 | 0.52303 | 0.52564 | -0.38916 | 0.086095 | 0.017237 | [9 empty] | 0-1 | 1.0438 | 0.54374 | 0.52038 | -0.08779 | -0.69036 | 0.056107 | 0-1 | 1.0445 | 0.53339 | 0.52301 | -0.23847 | -0.30213 | 0.036672 | 0-1 |  |  |  | 0.53062 | 0.23759 | 0.054754
1-2 | 1.0388 | 0.55012 | 0.53204 | -0.32257 | -0.081457 | 0.013584 | MEASURE | Lat (◦) | uncertainty associated to gravity in situ |  |  |  | ALTITUDE | MEASURE |  | 1-2 | 1.0243 | 0.54632 | 0.54228 | -0.17685 | -0.033717 | 0.076601 | 1-2 | 1.0315 | 0.54822 | 0.53716 | -0.24971 | -0.057587 | 0.045092 | 1-2 |  |  |  | 0.49257 | 0.28281 | 0.065278
2-3 | 1.0479 | 0.53561 | 0.52719 | -0.2862 | 0.024065 | 0.054996 | 1 | 0.17356 |  |  |  |  | 1225.96 | 1 |  | 2-3 | 1.0568 | 0.57195 | 0.5275 | -0.25761 | -0.10508 | 0.051157 | 2-3 | 1.0523 | 0.55378 | 0.52735 | -0.27191 | -0.040509 | 0.053076 | 2-3 |  |  |  | 0.49631 | 0.23738 | 0.061375
3-4 | 1.0604 | 0.59774 | 0.54364 | -0.22487 | -0.043233 | 0.023285 | 2 | 0.17356 | u(g) | 0.00056454 |  |  | 1226.073 | 2 |  | 3-4 | 1.0491 | 0.58253 | 0.54422 | -0.27317 | -0.31646 | 0.08826 | 3-4 | 1.0547 | 0.59014 | 0.54393 | -0.24902 | -0.17985 | 0.055772 | 3-4 |  |  |  | 0.47327 | 0.27643 | 0.067606
4-5 | 1.0548 | 0.56282 | 0.54587 | -0.23349 | 0.14761 | 0.013319 | 3 | 0.17355 | u(g(theta))^2 | 1.7716e-07 |  |  | 1225.907 | 3 |  | 4-5 | 1.0472 | 0.57175 | 0.52486 | -0.26338 | -0.16269 | 0.085532 | 4-5 | 1.051 | 0.56728 | 0.53536 | -0.24843 | -0.0075382 | 0.049426 | 4-5 |  |  |  | 0.47393 | 0.26166 | 0.055875
5-10 | 1.0567 | 0.56803 | 0.53965 | -0.27541 | -0.033513 | 0.015959 | 4 | 0.17358 | u(g(H)) | 1.4788e-07 |  |  | 1225.918 | 4 |  | 5-10 | 1.0482 | 0.57952 | 0.54038 | -0.34451 | -0.19074 | 0.074735 | 5-10 | 1.0524 | 0.57377 | 0.54002 | -0.30996 | -0.11213 | 0.045347 | 5-10 |  |  |  | 0.48022 | 0.25977 | 0.05117
10-15 | 1.0539 | 0.56065 | 0.53902 | -0.39725 | -0.17957 | 0.024214 | 5 | 0.17355 | uform | 0.00056454 |  |  | 1226.048 | 5 |  | 10-15 | 1.0282 | 0.592 | 0.53905 | -0.38143 | -0.24019 | 0.051743 | 10-15 | 1.041 | 0.57632 | 0.53904 | -0.38934 | -0.20988 | 0.037979 | 10-15 |  |  |  | 0.46561 | 0.22993 | 0.045864
15-20 | 1.042 | 0.57939 | 0.54028 | -0.24524 | -0.10171 | 0.034586 | VD | 1e-07 |  |  |  |  | 0.001 | VD |  | 15-20 | 1.0396 | 0.57755 | 0.53473 | -0.29308 | -0.24707 | 0.07207 | 15-20 | 1.0408 | 0.57847 | 0.53751 | -0.26916 | -0.17439 | 0.053328 | 15-20 |  |  |  | 0.47936 | 0.20854 | 0.051807
20-25 | 1.1379 | 0.62116 | 0.54003 | -0.59178 | -0.19624 | 0.040561 | u(VD) | 2.8868e-08 |  |  |  |  | 0.00028868 | u(VD) |  | 20-25 | 1.039 | 0.59202 | 0.5577 | -0.36892 | -0.2153 | 0.088316 | 20-25 | 1.0885 | 0.60659 | 0.54887 | -0.48035 | -0.20577 | 0.064439 | 20-25 |  |  |  | 0.50043 | 0.23631 | 0.055642
25-30 | 1.0625 | 0.58783 | 0.53961 | -0.37121 | 0.10248 | 0.062403 | urep | 1.0084e-05 |  |  |  |  | 0.04792 | urep |  | 25-30 | 1.0457 | 0.60236 | 0.5342 | -0.4846 | -0.27636 | 0.079911 | 25-30 | 1.0541 | 0.5951 | 0.5369 | -0.42791 | -0.08694 | 0.071157 | 25-30 |  |  |  | 0.49072 | 0.231 | 0.056197
30-35 | 1.0548 | 0.57921 | 0.54234 | -0.20564 | -0.093136 | 0.031927 | Ucomb | 1.0084e-05 |  |  |  |  | 0.047921 | Ucomb |  | 30-35 | 1.0417 | 0.59379 | 0.5391 | -0.46554 | -0.29564 | 0.10099 | 30-35 | 1.0483 | 0.5865 | 0.54072 | -0.33559 | -0.19439 | 0.066458 | 30-35 |  |  |  | 0.50605 | 0.29326 | 0.056227
35-40 | 1.0722 | 0.58589 | 0.54446 | -0.26115 | -0.081658 | 0.039426 | Lat(prom) | 0.17356 |  |  |  |  | 1225.9812 | H(prom) |  | 35-40 | 1.0392 | 0.67093 | 0.53264 | -0.48774 | -0.14279 | 0.12446 | 35-40 | 1.0557 | 0.62841 | 0.53855 | -0.37444 | -0.11222 | 0.081942 | 35-40 |  |  |  | 0.52315 | 0.24153 | 0.060631
40-45 | 1.0437 | 0.5989 | 0.53879 | -0.59269 | -0.24966 | 0.022807 | [9 empty] | 40-45 | 1.0365 | 0.65405 | 0.54129 | -0.54325 | -0.39364 | 0.068696 | 40-45 | 1.0401 | 0.62648 | 0.54004 | -0.56797 | -0.32165 | 0.045751 | 40-45 |  |  |  | 0.5205 | 0.24548 | 0.054642
45-50 | 1.0533 | 0.57771 | 0.53754 | -0.43235 | -0.04909 | 0.033279 | [9 empty] | 45-50 | 1.0502 | 0.62473 | 0.52457 | -0.5082 | -0.41096 | 0.088976 | 45-50 | 1.0517 | 0.60122 | 0.53105 | -0.47028 | -0.23002 | 0.061127 | 45-50 |  |  |  | 0.53033 | 0.20551 | 0.054701
50-55 | 1.0704 | 0.5856 | 0.53222 | -0.37335 | -0.071538 | 0.011874 | [9 empty] | 50-55 | 1.0398 | 0.59914 | 0.5293 | -0.48427 | -0.33836 | 0.092529 | 50-55 | 1.0551 | 0.59237 | 0.53076 | -0.42881 | -0.20495 | 0.052202 | 50-55 |  |  |  | 0.53351 | 0.2484 | 0.05522 |  |  | k | 2
55-60 | 1.074 | 0.54257 | 0.52326 | -0.4463 | 0.0088214 | 0.0057952 | [9 empty] | 55-60 | 1.0451 | 0.61187 | 0.53828 | -0.47692 | -0.36191 | 0.10928 | 55-60 | 1.0595 | 0.57722 | 0.53077 | -0.46161 | -0.17654 | 0.057535 | 55-60 |  |  |  | 0.55727 | 0.22062 | 0.057616
[23 empty] | IL (INDICATION WITH LOAD) | [6 empty] | Combined uncertainty for indication | [8 empty] | Exp uncertainty indic
40 kg -- REP UNIQUE | [6 empty] | UTILIZED WEIGHTS |  | CODE | [6 empty] | 40 kg -- REP UNIQUE | [6 empty] | 40 kg -- REP UNIQUE | [6 empty] | 40 kg -- REP UNIQUE | [8 empty] | 40 kg -- REP unique |  |  | 40 kg -- REP UNIQUE
Intervalo (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg) | left foot |  | M-015 | [6 empty] | Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg) | Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg) | Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg) | n | veff | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg) | Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg)
0-1 | 1.052 | 0.55739 | 0.53846 | 391.8361 | -1.4634 | -1.3784 | MC |  | 195.5579 | ∑ | [5 empty] | 0-1 | 1.0351 | 0.55494 | 0.53537 | 392.2276 | -0.9711 | -1.2967 | 0-1 | 1.0435 | 0.55617 | 0.53691 | 0.9164 | -1.2172 | -1.3376 | 0-1 |  |  |  | 0.53837 | 0.25283 | 0.10232 | 1004 | 1071.7246 | 1.0767 | 0.50567 | 0.20463 | 0-1 | -0.00099626 | 0.022781 | 0.013904 | 1.1549 | -0.91511 | -1.3742
1-2 | 1.1055 | 0.57514 | 0.53008 | 392.0125 | -1.4854 | -1.3688 |  |  |  | 391.1154 | [5 empty] | 1-2 | 1.0123 | 0.56847 | 0.55111 | 392.2168 | -1.3215 | -1.2853 | 1-2 | 1.0589 | 0.5718 | 0.5406 | 0.99919 | -1.4035 | -1.327 | 1-2 |  |  |  | 0.50098 | 0.29571 | 0.10825 | 1004 | 1083.3799 | 1.002 | 0.59142 | 0.2165 | 1-2 | 0.027371 | 0.023582 | 0.0034352 | 1.2489 | -1.3459 | -1.3721
2-3 | 1.0244 | 0.54203 | 0.54153 | 391.9837 | -1.4391 | -1.3354 | right foot |  | M-022 | ∑ | [5 empty] | 2-3 | 1.0312 | 0.54702 | 0.53352 | 392.2846 | -1.1795 | -1.3051 | 2-3 | 1.0278 | 0.54452 | 0.53752 | 1.0187 | -1.3093 | -1.3202 | 2-3 |  |  |  | 0.50462 | 0.25268 | 0.10597 | 1004 | 1082.0926 | 1.0092 | 0.50535 | 0.21193 | 2-3 | -0.024522 | -0.0092586 | 0.010178 | 1.2906 | -1.2688 | -1.3733
3-4 | 1.0728 | 0.53342 | 0.5461 | 392.0814 | -1.6798 | -1.3594 | MC |  | 195.5575 | [6 empty] | 3-4 | 1.0664 | 0.558 | 0.53729 | 392.3695 | -0.94038 | -1.3143 | 3-4 | 1.0696 | 0.54571 | 0.54169 | 1.11 | -1.3101 | -1.3368 | 3-4 |  |  |  | 0.48202 | 0.28973 | 0.10972 | 1004 | 1090.5463 | 0.96404 | 0.57947 | 0.21943 | 3-4 | 0.014887 | -0.04443 | -0.0022378 | 1.359 | -1.1303 | -1.3926
4-5 | 1.0639 | 0.55091 | 0.52348 | 392.107 | -1.4769 | -1.363 | [9 empty] | 4-5 | 1.02 | 0.55204 | 0.53819 | 392.4268 | -1.0568 | -1.3278 | 4-5 | 1.0419 | 0.55147 | 0.53084 | 1.1515 | -1.2669 | -1.3454 | 4-5 |  |  |  | 0.48263 | 0.27555 | 0.10287 | 1004 | 1089.546 | 0.96526 | 0.55111 | 0.20575 | 4-5 | -0.0090859 | -0.015811 | -0.0045237 | 1.3999 | -1.2593 | -1.3948
5-10 | 1.0424 | 0.55992 | 0.5325 | 392.2775 | -1.537 | -1.3465 | [9 empty] | 5-10 | 1.0489 | 0.64595 | 0.52544 | 392.5442 | -1.1889 | -1.3168 | 5-10 | 1.0456 | 0.60294 | 0.52897 | 1.2954 | -1.3629 | -1.3317 | 5-10 |  |  |  | 0.48787 | 0.27334 | 0.099232 | 5020 | 5356.9101 | 0.97573 | 0.54668 | 0.19846 | 5-10 | -0.0067945 | 0.029161 | -0.011045 | 1.6054 | -1.2508 | -1.377
10-15 | 1.0405 | 0.59513 | 0.5336 | 392.3722 | -1.4287 | -1.3614 | Uncertainty OFFSET WEIGHTS@40 kg | [8 empty] | 10-15 | 1.0441 | 0.64462 | 0.53557 | 392.5893 | -1.1892 | -1.2965 | 10-15 | 1.0423 | 0.61987 | 0.53459 | 1.3653 | -1.309 | -1.329 | 10-15 |  |  |  | 0.47349 | 0.24517 | 0.096606 | 5021 | 5379.5072 | 0.94697 | 0.49033 | 0.19321 | 10-15 | 0.0012949 | 0.043549 | -0.0044529 | 1.7546 | -1.0991 | -1.3669
15-20 | 1.0298 | 0.56715 | 0.53463 | 392.6307 | -1.4979 | -1.3986 | u (x) | 0.00025981 |  | N | [5 empty] | 15-20 | 1.0389 | 0.62126 | 0.52913 | 392.6932 | -1.1165 | -1.318 | 15-20 | 1.0344 | 0.5942 | 0.53188 | 1.5465 | -1.3072 | -1.3583 | 15-20 |  |  |  | 0.48702 | 0.22522 | 0.099568 | 5020 | 5358.2261 | 0.97404 | 0.45043 | 0.19914 | 15-20 | -0.0064368 | 0.015734 | -0.0056261 | 1.8157 | -1.1328 | -1.4117
20-25 | 1.0528 | 0.57881 | 0.53419 | 392.7329 | -1.5231 | -1.4271 | k | 3 | [7 empty] | 20-25 | 1.0434 | 0.65307 | 0.52624 | 392.7756 | -1.0452 | -1.3244 | 20-25 | 1.0481 | 0.61594 | 0.53022 | 1.6388 | -1.2841 | -1.3758 | 20-25 |  |  |  | 0.50777 | 0.25116 | 0.10161 | 5020 | 5329.8898 | 1.0155 | 0.50231 | 0.20323 | 20-25 | -0.04036 | 0.009348 | -0.018646 | 2.1192 | -1.0784 | -1.4402
25-30 | 1.0543 | 0.58552 | 0.53297 | 392.8319 | -1.362 | -1.4211 | Uncertainty OFFSET WEIGHTS@40 kg | [8 empty] | 25-30 | 1.0275 | 0.6209 | 0.53651 | 392.8732 | -1.1382 | -1.3309 | 25-30 | 1.0409 | 0.60321 | 0.53474 | 1.7371 | -1.2501 | -1.376 | 25-30 |  |  |  | 0.4982 | 0.24617 | 0.10192 | 5020 | 5342.4603 | 0.9964 | 0.49233 | 0.20385 | 25-30 | -0.013246 | 0.0081165 | -0.0021623 | 2.165 | -1.1632 | -1.4472
30-35 | 1.0523 | 0.56124 | 0.53717 | 392.8932 | -1.4851 | -1.4352 | u (x) | 7.5e-05 |  | N | [5 empty] | 30-35 | 1.0562 | 0.55736 | 0.53234 | 392.9042 | -1.1225 | -1.3121 | 30-35 | 1.0543 | 0.5593 | 0.53475 | 1.7833 | -1.3038 | -1.3736 | 30-35 |  |  |  | 0.51331 | 0.30534 | 0.10193 | 5020 | 5322.8834 | 1.0266 | 0.61068 | 0.20387 | 30-35 | 0.0060062 | -0.027197 | -0.0059673 | 2.1189 | -1.1094 | -1.4401
35-40 | 1.0244 | 0.54855 | 0.54222 | 392.8991 | -1.4453 | -1.4052 | k | 2 | [7 empty] | 35-40 | 1.0475 | 0.59274 | 0.52656 | 392.9804 | -1.0363 | -1.3086 | 35-40 | 1.0359 | 0.57064 | 0.53439 | 1.8243 | -1.2408 | -1.3569 | 35-40 |  |  |  | 0.53018 | 0.25607 | 0.10443 | 5020 | 5303.0117 | 1.0604 | 0.51215 | 0.20886 | 35-40 | -0.019781 | -0.057769 | -0.004159 | 2.1988 | -1.1286 | -1.4389
40-45 | 1.0392 | 0.5519 | 0.53229 | 392.993 | -1.3556 | -1.4325 | Uncertainty OFFSET WEIGHTS@40 kg | [8 empty] | 40-45 | 1.0253 | 0.58265 | 0.53884 | 392.9429 | -1.1011 | -1.3033 | 40-45 | 1.0322 | 0.56728 | 0.53556 | 1.8525 | -1.2284 | -1.3679 | 40-45 |  |  |  | 0.52755 | 0.2598 | 0.10107 | 5020 | 5305.78 | 1.0551 | 0.5196 | 0.20215 | 40-45 | -0.0078435 | -0.059198 | -0.0044731 | 2.4205 | -0.90672 | -1.4136
45-50 | 1.0318 | 0.54843 | 0.53789 | 393.0532 | -1.4723 | -1.4609 | u (x) | 0.084681 |  | N | [5 empty] | 45-50 | 1.0332 | 0.56505 | 0.53074 | 392.8305 | -1.2886 | -1.3217 | 45-50 | 1.0325 | 0.55674 | 0.53431 | 1.8264 | -1.3805 | -1.3913 | 45-50 |  |  |  | 0.53727 | 0.22242 | 0.10111 | 5020 | 5295.1984 | 1.0745 | 0.44485 | 0.20222 | 45-50 | -0.019287 | -0.044479 | 0.0032605 | 2.2967 | -1.1504 | -1.4524
50-55 | 1.0437 | 0.59131 | 0.53579 | 393.0893 | -1.5044 | -1.4549 | [9 empty] | 50-55 | 1.0273 | 0.62559 | 0.53193 | 392.9114 | -1.1658 | -1.3043 | 50-55 | 1.0355 | 0.60845 | 0.53386 | 1.8849 | -1.3351 | -1.3796 | 50-55 |  |  |  | 0.54041 | 0.26256 | 0.10139 | 5020 | 5291.9218 | 1.0808 | 0.52512 | 0.20277 | 50-55 | -0.019611 | 0.016082 | 0.0031055 | 2.3137 | -1.1301 | -1.4318
55-60 | 1.0377 | 0.57523 | 0.53176 | 393.1312 | -1.4656 | -1.4632 | [9 empty] | 55-60 | 1.0366 | 0.61396 | 0.53432 | 392.9232 | -1.115 | -1.313 | 55-60 | 1.0371 | 0.5946 | 0.53304 | 1.9118 | -1.2903 | -1.3881 | 55-60 |  |  |  | 0.56388 | 0.23646 | 0.10272 | 4780 | 5017.5459 | 1.1278 | 0.47292 | 0.20545 | 55-60 | -0.022401 | 0.017377 | 0.0022664 | 2.3734 | -1.1138 | -1.4456
80 kg -- REP unique | [6 empty] | UTILIZED WEIGHTS |  | CODE | CODE | [5 empty] | 80 kg -- REP unique | [6 empty] | 80 kg -- REP unique | [6 empty] | 80 kg -- REP unique | [8 empty] | 80 kg -- REP unique |  |  | 80 kg -- REP unique
Intervalo (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg) | left foot |  | M-015 | M-016 | ∑ |  |  |  |  | Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg) | Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg) | Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg) | n | veff | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg) | Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg)
0-1 | 1.0258 | 0.53654 | 0.54136 | 784.774 | -2.9119 | -2.7061 | MC |  | 195.5579 | 195.5534 | 391.1113 | ∑ |  |  |  | 0-1 | 1.0252 | 0.54301 | 0.54319 | 784.3816 | -2.8149 | -2.6601 | 0-1 | 1.0255 | 0.53978 | 0.54227 | 2.3541 | -2.8634 | -2.6831 | 0-1 |  |  |  | 0.55816 | 0.29262 | 0.17937 | 1004 | 1071.7246 | 1.1163 | 0.58525 | 0.35874 | 0-1 | -0.019 | 0.0063927 | 0.019262 | 2.5926 | -2.5613 | -2.7198 | Exp u | Interval (s) | Indication error (N) | HIGH ±u(z) (N) | LOW ±u(z) (N) | Close | diff
1-2 | 1.0429 | 0.51402 | 0.5359 | 784.7119 | -3.1544 | -2.6956 | [5 empty] | 782.2237 |  |  |  | 1-2 | 1.0381 | 0.53378 | 0.5649 | 784.6162 | -2.9787 | -2.6742 | 1-2 | 1.0405 | 0.5239 | 0.5504 | 2.4404 | -3.0666 | -2.6849 | 1-2 |  |  |  | 0.5222 | 0.33038 | 0.18282 | 1004 | 1083.3799 | 1.0444 | 0.66076 | 0.36563 | 1-2 | 0.0089627 | -0.024319 | 0.013241 | 2.6901 | -3.009 | -2.73 | 1.1163 | 0-1 | 2.5926 | 3.7089 | 1.4763 | 2.5926 | 2.2327
2-3 | 1.0533 | 0.54281 | 0.52374 | 784.6908 | -3.0462 | -2.6603 | right foot |  | M-022 | M-018 | ∑ | ∑ |  |  |  | 2-3 | 1.0318 | 0.54188 | 0.55093 | 784.6873 | -2.8238 | -2.6874 | 2-3 | 1.0426 | 0.54235 | 0.53734 | 2.4653 | -2.935 | -2.6739 | 2-3 |  |  |  | 0.52569 | 0.29249 | 0.18147 | 1004 | 1082.0926 | 1.0514 | 0.58498 | 0.36295 | 2-3 | -0.0097143 | -0.011434 | 0.0099901 | 2.7372 | -2.8945 | -2.7269 | 1.0444 | 1-2 | 2.6901 | 3.7345 | 1.6457 | 2.6901 | 2.0888
3-4 | 1.0221 | 0.55818 | 0.5352 | 784.8644 | -3.0701 | -2.6855 | MC |  | 195.5575 | 195.5549 | 391.1124 |  |  |  |  | 3-4 | 1.1208 | 0.52895 | 0.55011 | 784.6904 | -2.7457 | -2.6983 | 3-4 | 1.0714 | 0.54356 | 0.54266 | 2.5537 | -2.9079 | -2.6919 | 3-4 |  |  |  | 0.50403 | 0.32504 | 0.18369 | 1004 | 1090.5463 | 1.0081 | 0.65008 | 0.36738 | 3-4 | 0.016693 | -0.046574 | -0.0012706 | 2.8028 | -2.7281 | -2.7477 | 1.0514 | 2-3 | 2.7372 | 3.7886 | 1.6859 | 2.7372 | 2.1028
4-5 | 1.0394 | 0.6213 | 0.53765 | 784.9897 | -3.258 | -2.7226 | [9 empty] | 4-5 | 1.0389 | 0.55186 | 0.53548 | 784.7991 | -2.7779 | -2.6857 | 4-5 | 1.0392 | 0.58658 | 0.53657 | 2.6708 | -3.0179 | -2.7041 | 4-5 |  |  |  | 0.50461 | 0.31247 | 0.17969 | 1004 | 1089.546 | 1.0092 | 0.62493 | 0.35937 | 4-5 | -0.011834 | 0.019294 | 0.0012059 | 2.9192 | -3.0104 | -2.7535 | 1.0081 | 3-4 | 2.8028 | 3.8108 | 1.7947 | 2.8028 | 2.0161
5-10 | 1.0707 | 0.5451 | 0.52829 | 785.1673 | -2.9824 | -2.6949 | Uncertainty OFFSET WEIGHTS@80 kg | [8 empty] | 5-10 | 1.0945 | 0.54352 | 0.54889 | 784.888 | -2.7625 | -2.6838 | 5-10 | 1.0826 | 0.54431 | 0.53859 | 2.804 | -2.8725 | -2.6894 | 5-10 |  |  |  | 0.50962 | 0.31051 | 0.17763 | 5020 | 5356.9101 | 1.0192 | 0.62102 | 0.35525 | 5-10 | 0.030145 | -0.029466 | -0.0014225 | 3.114 | -2.7603 | -2.7347 | 1.0092 | 4-5 | 2.9192 | 3.9284 | 1.91 | 2.9192 | 2.0185
10-15 | 1.0603 | 0.59911 | 0.53681 | 785.5019 | -3.054 | -2.7139 | u (x) | 0.00052135 |  | N | [5 empty] | 10-15 | 1.1116 | 0.58548 | 0.53738 | 784.9905 | -2.809 | -2.6822 | 10-15 | 1.0859 | 0.5923 | 0.53709 | 3.0225 | -2.9315 | -2.698 | 10-15 |  |  |  | 0.49588 | 0.28603 | 0.17617 | 5021 | 5379.5072 | 0.99175 | 0.57205 | 0.35235 | 10-15 | 0.044902 | 0.015975 | -0.0019439 | 3.4119 | -2.7217 | -2.736 | 1.0192 | 5-10 | 3.114 | 4.1332 | 2.0947 | 3.114 | 2.0385
15-20 | 1.0409 | 0.58473 | 0.53471 | 785.5415 | -3.1304 | -2.713 | k | 3 | [7 empty] | 15-20 | 1.0388 | 0.58394 | 0.52492 | 785.2514 | -2.84 | -2.673 | 15-20 | 1.0399 | 0.58433 | 0.52981 | 3.1727 | -2.9852 | -2.693 | 15-20 |  |  |  | 0.50881 | 0.26912 | 0.17781 | 5020 | 5358.2261 | 1.0176 | 0.53824 | 0.35563 | 15-20 | -0.00092395 | 0.0058638 | -0.0076972 | 3.4419 | -2.8108 | -2.7463 | 0.99175 | 10-15 | 3.4119 | 4.4036 | 2.4201 | 3.4119 | 1.9835
20-25 | 1.0522 | 0.57977 | 0.54041 | 785.6678 | -2.9767 | -2.7636 | Uncertainty OFFSET WEIGHTS@80 kg | [8 empty] | 20-25 | 1.0488 | 0.58024 | 0.53384 | 785.2279 | -2.6697 | -2.6576 | 20-25 | 1.0505 | 0.58 | 0.53713 | 3.2242 | -2.8232 | -2.7106 | 20-25 |  |  |  | 0.52871 | 0.29118 | 0.17897 | 5020 | 5329.8898 | 1.0574 | 0.58235 | 0.35794 | 20-25 | -0.037944 | -0.026585 | -0.01174 | 3.7045 | -2.6174 | -2.775 | 1.0176 | 15-20 | 3.4419 | 4.4595 | 2.4243 | 3.4419 | 2.0353
25-30 | 1.0525 | 0.58052 | 0.54215 | 785.6738 | -3.0573 | -2.7427 | u (x) | 0.0001505 |  | N | [5 empty] | 25-30 | 1.0382 | 0.55996 | 0.54367 | 785.3288 | -2.409 | -2.6437 | 25-30 | 1.0453 | 0.57024 | 0.54291 | 3.2776 | -2.7331 | -2.6932 | 25-30 |  |  |  | 0.51953 | 0.28688 | 0.17914 | 5020 | 5342.4603 | 1.0391 | 0.57377 | 0.35829 | 25-30 | -0.0087947 | -0.024861 | 0.0060095 | 3.7055 | -2.6462 | -2.7644 | 1.0574 | 20-25 | 3.7045 | 4.7619 | 2.6471 | 3.7045 | 2.1148
30-35 | 1.0366 | 0.56363 | 0.53303 | 785.7774 | -3.0808 | -2.7577 | k | 2 | [7 empty] | 30-35 | 1.0253 | 0.61515 | 0.5368 | 785.5219 | -2.5671 | -2.6421 | 30-35 | 1.0309 | 0.58939 | 0.53491 | 3.426 | -2.824 | -2.6999 | 30-35 |  |  |  | 0.53403 | 0.33903 | 0.17915 | 5020 | 5322.8834 | 1.0681 | 0.67805 | 0.3583 | 30-35 | -0.017322 | 0.0028923 | -0.0058072 | 3.7616 | -2.6296 | -2.7663 | 1.0391 | 25-30 | 3.7055 | 4.7446 | 2.6665 | 3.7055 | 2.0781
35-40 | 1.0501 | 0.56132 | 0.54008 | 785.8539 | -3.1292 | -2.7797 | Uncertainty OFFSET WEIGHTS@80 kg | [8 empty] | 35-40 | 1.03 | 0.6112 | 0.53295 | 785.616 | -2.512 | -2.6411 | 35-40 | 1.0401 | 0.58626 | 0.53652 | 3.5112 | -2.8206 | -2.7104 | 35-40 |  |  |  | 0.55027 | 0.29543 | 0.18058 | 5020 | 5303.0117 | 1.1005 | 0.59086 | 0.36116 | 35-40 | -0.015647 | -0.04215 | -0.0020338 | 3.8857 | -2.7083 | -2.7923 | 1.0681 | 30-35 | 3.7616 | 4.8296 | 2.6935 | 3.7616 | 2.1361
40-45 | 1.0314 | 0.60988 | 0.53775 | 786.0658 | -2.8852 | -2.7743 | u (x) | 0.16993 |  | N | [5 empty] | 40-45 | 1.0324 | 0.60629 | 0.535 | 785.6182 | -2.7636 | -2.6639 | 40-45 | 1.0319 | 0.60808 | 0.53637 | 3.6183 | -2.8244 | -2.7191 | 40-45 |  |  |  | 0.54774 | 0.29866 | 0.17866 | 5020 | 5305.78 | 1.0955 | 0.59732 | 0.35732 | 40-45 | -0.0081983 | -0.018393 | -0.0036653 | 4.1863 | -2.5028 | -2.7649 | 1.1005 | 35-40 | 3.8857 | 4.9862 | 2.7852 | 3.8857 | 2.2011
45-50 | 1.0381 | 0.56375 | 0.53381 | 786.1578 | -2.9899 | -2.7734 | [9 empty] | 45-50 | 1.0505 | 0.62658 | 0.53928 | 785.6365 | -2.9034 | -2.6717 | 45-50 | 1.0443 | 0.59517 | 0.53654 | 3.6735 | -2.9466 | -2.7225 | 45-50 |  |  |  | 0.5571 | 0.26679 | 0.17868 | 5020 | 5295.1984 | 1.1142 | 0.53358 | 0.35736 | 45-50 | -0.0074152 | -0.0060545 | 0.00549 | 4.1437 | -2.7166 | -2.7836 | 1.0955 | 40-45 | 4.1863 | 5.2818 | 3.0908 | 4.1863 | 2.1909
50-55 | 1.0412 | 0.56589 | 0.53388 | 785.9211 | -3.0374 | -2.7879 | [9 empty] | 50-55 | 1.0255 | 0.58034 | 0.53825 | 785.7435 | -2.7387 | -2.6565 | 50-55 | 1.0334 | 0.57312 | 0.53607 | 3.6086 | -2.888 | -2.7222 | 50-55 |  |  |  | 0.56013 | 0.30107 | 0.17884 | 5020 | 5291.9218 | 1.1203 | 0.60214 | 0.35768 | 50-55 | -0.021746 | -0.019254 | 0.0053087 | 4.0374 | -2.6831 | -2.7744 | 1.1142 | 45-50 | 4.1437 | 5.2579 | 3.0295 | 4.1437 | 2.2284
55-60 | 1.0699 | 0.58175 | 0.5332 | 786.1217 | -3.1347 | -2.7957 | [9 empty] | 55-60 | 1.0262 | 0.57234 | 0.52509 | 785.7784 | -2.8583 | -2.6794 | 55-60 | 1.048 | 0.57705 | 0.52914 | 3.7264 | -2.9965 | -2.7375 | 55-60 |  |  |  | 0.58281 | 0.2786 | 0.1796 | 4780 | 5017.5459 | 1.1656 | 0.5572 | 0.3592 | 55-60 | -0.011499 | -0.00017571 | -0.0016286 | 4.188 | -2.82 | -2.795 | 1.1203 | 50-55 | 4.0374 | 5.1577 | 2.9172 | 4.0374 | 2.2405
120 kg -- REP unique | [6 empty] | UTILIZED WEIGHTS |  | CODE | CODE | CODE |  |  |  |  | 120 kg -- REP unique | [6 empty] | 120 kg -- REP unique | [6 empty] | 120 kg -- REP unique | [8 empty] | 120 kg -- REP unique |  |  | 120 kg -- REP unique
Intervalo (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg) | left foot |  | M-015 | M-016 | M-019 | ∑ |  |  |  | Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg) | Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg) | Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg) | n | veff | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg) | Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg)
0-1 | 1.0257 | 0.52669 | 0.53988 | 1177.1287 | -4.5749 | -4.0484 | MC |  | 195.5579 | 195.5534 | 195.5548 | 586.6661 | ∑ |  |  | 0-1 | 1.0452 | 0.53069 | 0.54055 | 1176.865 | -4.4081 | -4.048 | 0-1 | 1.0355 | 0.52869 | 0.54021 | 3.668 | -4.4915 | -4.0482 | 0-1 |  |  |  | 0.58949 | 0.34868 | 0.261 | 1004 | 1071.7246 | 1.179 | 0.69736 | 0.52201 | 0-1 | -0.0090399 | -0.0046916 | 0.017206 | 3.9065 | -4.1893 | -4.0849
1-2 | 1.0403 | 0.59304 | 0.52555 | 1177.284 | -4.4324 | -4.042 | [6 empty] | 1173.3288 |  |  | 1-2 | 1.0502 | 0.53949 | 0.5501 | 1177.025 | -4.0087 | -4.0338 | 1-2 | 1.0452 | 0.56627 | 0.53783 | 3.8257 | -4.2205 | -4.0379 | 1-2 |  |  |  | 0.55555 | 0.38092 | 0.26339 | 1004 | 1083.3799 | 1.1111 | 0.76184 | 0.52677 | 1-2 | 0.013711 | 0.018046 | 0.00066506 | 4.0754 | -4.1629 | -4.083
2-3 | 1.011 | 0.50807 | 0.51973 | 1177.3912 | -4.5406 | -4.0926 | right foot |  | M-022 | M-018 | M-020 | ∑ | ∑ |  |  | 2-3 | 1.0517 | 0.49978 | 0.52931 | 1177.1935 | -4.2346 | -4.0301 | 2-3 | 1.0313 | 0.50392 | 0.52452 | 3.9635 | -4.3876 | -4.0613 | 2-3 |  |  |  | 0.55884 | 0.34857 | 0.26246 | 1004 | 1082.0926 | 1.1177 | 0.69714 | 0.52491 | 2-3 | -0.020969 | -0.049858 | -0.002826 | 4.2354 | -4.3471 | -4.1144
3-4 | 1.086 | 0.51496 | 0.55035 | 1177.4079 | -4.4739 | -4.0879 | MC |  | 195.5575 | 195.5549 | 195.5503 | 586.6627 |  |  |  | 3-4 | 1.0529 | 0.54042 | 0.5405 | 1177.1633 | -4.13 | -4.0474 | 3-4 | 1.0694 | 0.52769 | 0.54542 | 3.9568 | -4.302 | -4.0677 | 3-4 |  |  |  | 0.53852 | 0.3763 | 0.26399 | 1004 | 1090.5463 | 1.077 | 0.7526 | 0.52799 | 3-4 | 0.01472 | -0.06245 | 0.0014919 | 4.2058 | -4.1221 | -4.1234
4-5 | 1.0131 | 0.56974 | 0.54665 | 1177.4796 | -4.513 | -4.0836 | [9 empty] | 4-5 | 1.0373 | 0.54445 | 0.55845 | 1177.2633 | -3.926 | -4.0582 | 4-5 | 1.0252 | 0.55709 | 0.55255 | 4.0426 | -4.2195 | -4.0709 | 4-5 |  |  |  | 0.53906 | 0.36549 | 0.26122 | 1004 | 1089.546 | 1.0781 | 0.73098 | 0.52245 | 4-5 | -0.025868 | -0.010189 | 0.017188 | 4.2911 | -4.212 | -4.1203
5-10 | 1.0696 | 0.55929 | 0.5416 | 1177.6957 | -4.4995 | -4.1027 | Uncertainty OFFSET WEIGHTS@120 kg | [8 empty] | 5-10 | 1.0611 | 0.57161 | 0.53551 | 1177.4155 | -4.1849 | -4.0625 | 5-10 | 1.0653 | 0.56545 | 0.53855 | 4.2268 | -4.3422 | -4.0826 | 5-10 |  |  |  | 0.54375 | 0.36382 | 0.25981 | 5020 | 5356.9101 | 1.0875 | 0.72765 | 0.51962 | 5-10 | 0.012904 | -0.0083271 | -0.0014622 | 4.5367 | -4.2301 | -4.1279
10-15 | 1.0369 | 0.58082 | 0.53827 | 1177.8604 | -4.521 | -4.1091 | u (x) | 0.00078115 |  | N | [5 empty] | 10-15 | 1.0627 | 0.57885 | 0.52356 | 1177.7 | -4.1856 | -4.1138 | 10-15 | 1.0498 | 0.57983 | 0.53091 | 4.4514 | -4.3533 | -4.1115 | 10-15 |  |  |  | 0.53089 | 0.34316 | 0.25882 | 5021 | 5379.5072 | 1.0618 | 0.68633 | 0.51764 | 10-15 | 0.0087508 | 0.003512 | -0.0081274 | 4.8407 | -4.1434 | -4.1495
15-20 | 1.0325 | 0.55519 | 0.53362 | 1178.0852 | -4.5938 | -4.1448 | k | 3 | [7 empty] | 15-20 | 1.1267 | 0.59991 | 0.55961 | 1177.9602 | -4.1076 | -4.0748 | 15-20 | 1.0796 | 0.57755 | 0.54662 | 4.6939 | -4.3507 | -4.1098 | 15-20 |  |  |  | 0.54299 | 0.32921 | 0.25994 | 5020 | 5358.2261 | 1.086 | 0.65841 | 0.51988 | 15-20 | 0.038816 | -0.0009193 | 0.0091076 | 4.963 | -4.1763 | -4.1631
20-25 | 1.044 | 0.55479 | 0.53155 | 1178.2471 | -4.3209 | -4.1301 | Uncertainty OFFSET WEIGHTS@120 kg | [8 empty] | 20-25 | 1.0839 | 0.57984 | 0.53566 | 1177.9006 | -4.0961 | -4.1131 | 20-25 | 1.0639 | 0.56732 | 0.53361 | 4.745 | -4.2085 | -4.1216 | 20-25 |  |  |  | 0.56168 | 0.34747 | 0.26073 | 5020 | 5329.8898 | 1.1234 | 0.69494 | 0.52146 | 20-25 | -0.024553 | -0.039272 | -0.015259 | 5.2254 | -4.0027 | -4.1861
25-30 | 1.0588 | 0.54283 | 0.5299 | 1178.3324 | -4.4738 | -4.137 | u (x) | 0.0002255 |  | N | [5 empty] | 25-30 | 1.0667 | 0.59938 | 0.54219 | 1178.0234 | -4.0856 | -4.1166 | 25-30 | 1.0628 | 0.57111 | 0.53605 | 4.8491 | -4.2797 | -4.1268 | 25-30 |  |  |  | 0.55305 | 0.34388 | 0.26085 | 5020 | 5342.4603 | 1.1061 | 0.68776 | 0.5217 | 25-30 | 0.0086534 | -0.023991 | -0.00085412 | 5.277 | -4.1928 | -4.198
30-35 | 1.0607 | 0.5669 | 0.52892 | 1178.4318 | -4.4921 | -4.1439 | k | 2 | [7 empty] | 30-35 | 1.0356 | 0.5775 | 0.53618 | 1178.1159 | -4.0772 | -4.1144 | 30-35 | 1.0482 | 0.5722 | 0.53255 | 4.945 | -4.2847 | -4.1291 | 30-35 |  |  |  | 0.56669 | 0.38844 | 0.26085 | 5020 | 5322.8834 | 1.1334 | 0.77689 | 0.52171 | 30-35 | -9.0093e-05 | -0.014294 | -0.0081718 | 5.2806 | -4.0903 | -4.1956
35-40 | 1.0298 | 0.59876 | 0.52971 | 1178.615 | -4.4421 | -4.1587 | Uncertainty OFFSET WEIGHTS@120 kg | [8 empty] | 35-40 | 1.0493 | 0.54959 | 0.53186 | 1178.2619 | -4.0514 | -4.1134 | 35-40 | 1.0395 | 0.57417 | 0.53079 | 5.1096 | -4.2468 | -4.1361 | 35-40 |  |  |  | 0.58202 | 0.35104 | 0.26184 | 5020 | 5303.0117 | 1.164 | 0.70208 | 0.52368 | 35-40 | -0.016158 | -0.05424 | -0.0077645 | 5.4841 | -4.1345 | -4.218
40-45 | 1.0454 | 0.56121 | 0.53706 | 1178.6032 | -4.3889 | -4.1623 | u (x) | 0.25461 |  | N | [5 empty] | 40-45 | 1.0418 | 0.57669 | 0.53762 | 1178.3788 | -4.061 | -4.1294 | 40-45 | 1.0436 | 0.56895 | 0.53734 | 5.1622 | -4.2249 | -4.1458 | 40-45 |  |  |  | 0.57963 | 0.35376 | 0.26052 | 5020 | 5305.7799 | 1.1593 | 0.70753 | 0.52104 | 40-45 | 0.0035004 | -0.057527 | -0.0026958 | 5.7301 | -3.9033 | -4.1916
45-50 | 1.0347 | 0.58162 | 0.53718 | 1178.7644 | -4.4207 | -4.1936 | [9 empty] | 45-50 | 1.0384 | 0.58662 | 0.53482 | 1178.4487 | -4.2 | -4.1027 | 45-50 | 1.0366 | 0.58412 | 0.536 | 5.2777 | -4.3104 | -4.1482 | 45-50 |  |  |  | 0.58848 | 0.3273 | 0.26053 | 5020 | 5295.1984 | 1.177 | 0.6546 | 0.52106 | 45-50 | -0.015174 | -0.017103 | 0.0049459 | 5.748 | -4.0803 | -4.2093
50-55 | 1.0371 | 0.56552 | 0.53298 | 1178.9692 | -4.4028 | -4.1996 | [9 empty] | 50-55 | 1.0304 | 0.57324 | 0.54197 | 1178.3933 | -4.1348 | -4.1212 | 50-55 | 1.0337 | 0.56938 | 0.53747 | 5.3525 | -4.2688 | -4.1604 | 50-55 |  |  |  | 0.59135 | 0.3558 | 0.26064 | 5020 | 5291.9218 | 1.1827 | 0.7116 | 0.52128 | 50-55 | -0.021375 | -0.022992 | 0.0067171 | 5.7813 | -4.0639 | -4.2126
55-60 | 1.0656 | 0.57768 | 0.54098 | 1178.8058 | -4.4703 | -4.2201 | [9 empty] | 55-60 | 1.053 | 0.57082 | 0.53209 | 1178.4358 | -3.9923 | -4.1223 | 55-60 | 1.0593 | 0.57425 | 0.53653 | 5.292 | -4.2313 | -4.1712 | 55-60 |  |  |  | 0.61288 | 0.337 | 0.26116 | 4780 | 5017.5459 | 1.2258 | 0.67399 | 0.52233 | 55-60 | -0.00025891 | -0.0029722 | 0.0057609 | 5.7536 | -4.0547 | -4.2287
140 kg -- REP unique | [6 empty] | UTILIZED WEIGHTS |  | CODE | CODE | CODE | CODE | CODE |  |  | 140 kg -- REP unique | [6 empty] | 140 kg -- REP unique | [6 empty] | 140 kg -- REP unique | [8 empty] | 140 kg -- REP unique |  |  | 140 kg -- REP unique
Intervalo (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg) | left foot |  | M-015 | M-016 | M-019 | M-005 | M-006 | ∑ |  | Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg) | Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg) | Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg) | n | veff | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg) | Interval (s) | STDEV Fz_average (kg) | STDEV Fy_average (kg) | STDEV Fx_average (kg) | AVERAGE Fz (kg) | AVERAGE Fy (kg) | AVERAGE Fx (kg)
0-1 | 1.0579 | 0.54138 | 0.56682 | 1373.6996 | -5.1553 | -4.7742 | MC |  | 195.5579 | 195.5534 | 195.5548 | 48.8904 | 48.8914 | 684.4479 | ∑ | 0-1 | 1.1119 | 0.62058 | 0.5276 | 1373.5029 | -4.5032 | -4.7248 | 0-1 | 1.0849 | 0.58098 | 0.54721 | 4.7079 | -4.8293 | -4.7495 | 0-1 |  |  |  | 0.6381 | 0.42573 | 0.35748 | 1004 | 1071.7246 | 1.2762 | 0.85147 | 0.71496 | 0-1 | 0.040386 | 0.047595 | 0.024203 | 4.9464 | -4.5271 | -4.7862
1-2 | 1.1371 | 0.55551 | 0.60428 | 1373.7778 | -4.9267 | -4.7736 | [8 empty] | 1368.8933 | 1-2 | 1.1451 | 0.52544 | 0.54287 | 1373.5408 | -4.9635 | -4.741 | 1-2 | 1.1411 | 0.54048 | 0.57358 | 4.766 | -4.9451 | -4.7573 | 1-2 |  |  |  | 0.60689 | 0.45251 | 0.35922 | 1004 | 1083.3799 | 1.2138 | 0.90503 | 0.71845 | 1-2 | 0.10958 | -0.0077438 | 0.036416 | 5.0157 | -4.8875 | -4.8024
2-3 | 1.0538 | 0.5284 | 0.54827 | 1373.878 | -4.8374 | -4.775 | right foot |  | M-022 | M-018 | M-020 | M-007 |  | ∑ | ∑ | 2-3 | 1.0644 | 0.52284 | 0.52004 | 1373.5493 | -4.9878 | -4.7308 | 2-3 | 1.0591 | 0.52562 | 0.53416 | 4.8203 | -4.9126 | -4.7529 | 2-3 |  |  |  | 0.60989 | 0.42564 | 0.35854 | 1004 | 1082.0926 | 1.2198 | 0.85128 | 0.71708 | 2-3 | 0.0067495 | -0.028163 | 0.0068094 | 5.0922 | -4.8721 | -4.806
3-4 | 1.1052 | 0.54928 | 0.52562 | 1373.9644 | -4.8939 | -4.7703 | MC |  | 195.5575 | 195.5549 | 195.5503 | 97.7827 |  | 684.4454 |  | 3-4 | 1.0448 | 0.55479 | 0.54865 | 1373.4934 | -4.8431 | -4.724 | 3-4 | 1.075 | 0.55204 | 0.53713 | 4.8356 | -4.8685 | -4.7472 | 3-4 |  |  |  | 0.59133 | 0.44863 | 0.35967 | 1004 | 1090.5463 | 1.1827 | 0.89726 | 0.71934 | 3-4 | 0.020272 | -0.038102 | -0.0067974 | 5.0846 | -4.6886 | -4.8029
4-5 | 1.0124 | 0.65076 | 0.53595 | 1373.9025 | -4.9202 | -4.7687 | [9 empty] | 4-5 | 1.0846 | 0.58519 | 0.54722 | 1373.5507 | -4.5926 | -4.7097 | 4-5 | 1.0485 | 0.61798 | 0.54158 | 4.8333 | -4.7564 | -4.7392 | 4-5 |  |  |  | 0.59182 | 0.43961 | 0.35764 | 1004 | 1089.546 | 1.1836 | 0.87921 | 0.71528 | 4-5 | -0.0025197 | 0.050693 | 0.0062215 | 5.0817 | -4.7488 | -4.7886
5-10 | 1.066 | 0.54506 | 0.5305 | 1373.9604 | -4.9836 | -4.7982 | Uncertainty OFFSET WEIGHTS@140 kg | [8 empty] | 5-10 | 1.0615 | 0.55523 | 0.53473 | 1373.6954 | -4.7957 | -4.7374 | 5-10 | 1.0637 | 0.55015 | 0.53261 | 4.9346 | -4.8896 | -4.7678 | 5-10 |  |  |  | 0.5961 | 0.43822 | 0.35661 | 5020 | 5356.9101 | 1.1922 | 0.87644 | 0.71322 | 5-10 | 0.011311 | -0.023629 | -0.0074031 | 5.2445 | -4.7775 | -4.8131
10-15 | 1.0502 | 0.58336 | 0.52731 | 1374.2242 | -4.756 | -4.8008 | u (x) | 0.0010825 |  | N | [5 empty] | 10-15 | 1.0799 | 0.54757 | 0.53465 | 1374.0843 | -4.8136 | -4.7372 | 10-15 | 1.065 | 0.56547 | 0.53098 | 5.2609 | -4.7848 | -4.769 | 10-15 |  |  |  | 0.58439 | 0.42122 | 0.35589 | 5021 | 5379.5072 | 1.1688 | 0.84245 | 0.71178 | 10-15 | 0.023965 | -0.010855 | -0.0080583 | 5.6503 | -4.5749 | -4.807
15-20 | 1.0627 | 0.54946 | 0.53275 | 1374.4166 | -4.9608 | -4.8292 | k | 3 | [7 empty] | 15-20 | 1.0542 | 0.57011 | 0.53776 | 1374.2478 | -4.665 | -4.7396 | 15-20 | 1.0585 | 0.55979 | 0.53526 | 5.4389 | -4.8129 | -4.7844 | 15-20 |  |  |  | 0.59541 | 0.40993 | 0.3567 | 5020 | 5358.2261 | 1.1908 | 0.81987 | 0.71341 | 15-20 | 0.017701 | -0.018684 | -0.0022538 | 5.7081 | -4.6385 | -4.8378
20-25 | 1.0539 | 0.53951 | 0.52945 | 1374.383 | -5.0496 | -4.8253 | Uncertainty OFFSET WEIGHTS@140 kg | [8 empty] | 20-25 | 1.0485 | 0.55989 | 0.53746 | 1374.452 | -4.7444 | -4.7416 | 20-25 | 1.0512 | 0.5497 | 0.53345 | 5.5242 | -4.897 | -4.7834 | 20-25 |  |  |  | 0.6125 | 0.42474 | 0.35728 | 5020 | 5329.8898 | 1.225 | 0.84948 | 0.71456 | 20-25 | -0.03729 | -0.056889 | -0.015414 | 6.0045 | -4.6912 | -4.8479
25-30 | 1.0463 | 0.57945 | 0.54355 | 1374.4861 | -4.7508 | -4.809 | u (x) | 0.0003125 |  | N | [5 empty] | 25-30 | 1.0346 | 0.56876 | 0.54094 | 1374.5133 | -4.6526 | -4.7511 | 25-30 | 1.0405 | 0.5741 | 0.54224 | 5.6064 | -4.7017 | -4.7801 | 25-30 |  |  |  | 0.60459 | 0.42181 | 0.35737 | 5020 | 5342.4603 | 1.2092 | 0.84361 | 0.71474 | 25-30 | -0.013667 | -0.020994 | 0.0053435 | 6.0343 | -4.6148 | -4.8512
30-35 | 1.0706 | 0.55882 | 0.54296 | 1374.659 | -4.9573 | -4.8685 | k | 2 | [7 empty] | 30-35 | 1.0551 | 0.58851 | 0.5392 | 1374.5528 | -4.6628 | -4.7738 | 30-35 | 1.0628 | 0.57366 | 0.54108 | 5.7126 | -4.8101 | -4.8212 | 30-35 |  |  |  | 0.6171 | 0.45887 | 0.35737 | 5020 | 5322.8834 | 1.2342 | 0.91773 | 0.71474 | 30-35 | 0.014561 | -0.012836 | 0.00035719 | 6.0482 | -4.6157 | -4.8876
35-40 | 1.046 | 0.60754 | 0.5394 | 1374.7452 | -4.8647 | -4.8436 | Uncertainty OFFSET WEIGHTS@140 kg | [8 empty] | 35-40 | 1.0506 | 0.54882 | 0.54591 | 1374.6254 | -4.6958 | -4.7578 | 35-40 | 1.0483 | 0.57818 | 0.54265 | 5.792 | -4.7802 | -4.8007 | 35-40 |  |  |  | 0.6312 | 0.42766 | 0.35809 | 5020 | 5303.0117 | 1.2624 | 0.85533 | 0.71618 | 35-40 | -0.0074541 | -0.050234 | 0.0041023 | 6.1664 | -4.668 | -4.8827
40-45 | 1.0674 | 0.55111 | 0.53624 | 1374.7818 | -4.9141 | -4.8485 | u (x) | 0.35284 |  | N | [5 empty] | 40-45 | 1.0463 | 0.57274 | 0.53661 | 1374.7584 | -4.7911 | -4.808 | 40-45 | 1.0569 | 0.56193 | 0.53642 | 5.8768 | -4.8526 | -4.8282 | 40-45 |  |  |  | 0.62899 | 0.42991 | 0.35713 | 5020 | 5305.7799 | 1.258 | 0.85981 | 0.71425 | 40-45 | 0.016785 | -0.06455 | -0.003615 | 6.4447 | -4.531 | -4.874
45-50 | 1.0663 | 0.55504 | 0.53468 | 1374.937 | -4.8217 | -4.8285 | [9 empty] | 45-50 | 1.0553 | 0.54718 | 0.53934 | 1374.7907 | -4.7182 | -4.7664 | 45-50 | 1.0608 | 0.55111 | 0.53701 | 5.9705 | -4.77 | -4.7974 | 45-50 |  |  |  | 0.63716 | 0.40841 | 0.35714 | 5020 | 5295.1984 | 1.2743 | 0.81681 | 0.71427 | 45-50 | 0.0090529 | -0.050115 | 0.0059581 | 6.4408 | -4.5399 | -4.8586
50-55 | 1.0492 | 0.54861 | 0.53824 | 1374.9405 | -4.9569 | -4.8621 | [9 empty] | 50-55 | 1.0656 | 0.57077 | 0.53447 | 1374.9524 | -4.7675 | -4.7509 | 50-55 | 1.0574 | 0.55969 | 0.53636 | 6.0531 | -4.8622 | -4.8065 | 50-55 |  |  |  | 0.63981 | 0.43158 | 0.35722 | 5020 | 5291.9218 | 1.2796 | 0.86316 | 0.71443 | 50-55 | 0.0022771 | -0.032679 | 0.0055999 | 6.482 | -4.6573 | -4.8587
55-60 | 1.0456 | 0.57449 | 0.53453 | 1375.1127 | -4.8236 | -4.8582 | [9 empty] | 55-60 | 1.0577 | 0.57497 | 0.53216 | 1374.9797 | -4.7239 | -4.7413 | 55-60 | 1.0516 | 0.57473 | 0.53335 | 6.1529 | -4.7737 | -4.7997 | 55-60 |  |  |  | 0.65976 | 0.41622 | 0.3576 | 4780 | 5017.5459 | 1.3195 | 0.83243 | 0.71519 | 55-60 | -0.0079066 | -0.0024882 | 0.0025738 | 6.6145 | -4.5972 | -4.8573
UNLOAD

## HysteresisDETERMINATION_ANOVA
2 WAY ANOVA
[8 empty] | FUENTE | gl | SC | CM | F | w2 |  | N
[7 empty] | A | 89032036.9083 | 15 | 20.5189 | 1.3679 | 9.2966 | 4.5244e-07 |  | #a | 16
| SOURCE | DF | SC | CM | F |  | B | 89032017.0814 | 1 | 0.69196 | 0.69196 | 2.9584 | 1.1318e-08 |  | #b | 2
A | EA^2/bs | (a-1) | A-T | SCA/glA | CMA/CMAS |  | AB | 89032037.7552 | 15 | 0.1549 | 0.010327 | 0.88562 | -4.9429e-10 |  | s | 5
B | EB^2/as | (b-1) | B-T | SCB/glB | CMB/CMBS |  | S | 129505336.8217 | 4 | 40473320.4323 | 10118330.1081
AB | E(AB)^2/s | (a-1)(b-1) | AB-B-A+T | SCAB/glAB | CMAB/CMABS |  | AS | 129505366.1692 | 60 | 8.8286 | 0.14714
S | Es^2/ab | (s-1) | S-T | SCS/glS |  |  | BS | 129505338.4492 | 4 | 0.93559 | 0.2339
AS | E(AS)^2/b | (a-1)(s-1) | AS-A-S+T | SCAS/glAS |  |  | ABS | 129505368.6513 | 60 | 0.69961 | 0.01166
BS | E(BS)^2/a | (b-1)(s-1) | BS-B-S+T | SCBS/glBS |  |  | T | 89032016.3894 | 159 | 40473352.2619 | 254549.3853
ABS | E(ABS)^2 | (a-1)(b-1)(s-1) | ABS-AB-BS-AS+A+B+S-T | SCABS/glABS
T | E(T)^2/abs | abs-1 | ABS-T | SCT/glT | [7 empty] | F(15,60) | 1.84
[9 empty] | p<0.05 |  |  | F(1,4) | 7.71
[12 empty] | F(15,60) | 1.84
A--> | 0-1 | 1-2 | 2-3 | 3-4 | 4-5 | 5-10 | 10-15 | 15-20 | 20-25 | 25-30 | 30-35 | 35-40 | 40-45 | 45-50 | 50-55 | 55-60 | Ss
LOAD=b1 | -0.38916 | -0.32257 | -0.2862 | -0.22487 | -0.23349 | -0.27541 | -0.39725 | -0.24524 | -0.59178 | -0.37121 | -0.20564 | -0.26115 | -0.59269 | -0.43235 | -0.37335 | -0.4463 | -5.6487
| 391.8361 | 392.0125 | 391.9837 | 392.0814 | 392.107 | 392.2775 | 392.3722 | 392.6307 | 392.7329 | 392.8319 | 392.8932 | 392.8991 | 392.993 | 393.0532 | 393.0893 | 393.1312 | 6280.9251
| 784.774 | 784.7119 | 784.6908 | 784.8644 | 784.9897 | 785.1673 | 785.5019 | 785.5415 | 785.6678 | 785.6738 | 785.7774 | 785.8539 | 786.0658 | 786.1578 | 785.9211 | 786.1217 | 12567.4808
| 1177.1287 | 1177.284 | 1177.3912 | 1177.4079 | 1177.4796 | 1177.6957 | 1177.8604 | 1178.0852 | 1178.2471 | 1178.3324 | 1178.4318 | 1178.615 | 1178.6032 | 1178.7644 | 1178.9692 | 1178.8058 | 18849.1015
| 1373.6996 | 1373.7778 | 1373.878 | 1373.9644 | 1373.9025 | 1373.9604 | 1374.2242 | 1374.4166 | 1374.383 | 1374.4861 | 1374.659 | 1374.7452 | 1374.7818 | 1374.937 | 1374.9405 | 1375.1127 | 21989.8689
AVERAGE | 745.4099 | 745.4927 | 745.5315 | 745.6186 | 745.6491 | 745.7651 | 745.9123 | 746.0857 | 746.0878 | 746.1906 | 746.3112 | 746.3704 | 746.3702 | 746.496 | 746.5094 | 746.545 | 11936.3455
Ex | 3727.0493 | 3727.4637 | 3727.6574 | 3728.0932 | 3728.2454 | 3728.8255 | 3729.5614 | 3730.4287 | 3730.439 | 3730.953 | 3731.5558 | 3731.8521 | 3731.8511 | 3732.48 | 3732.5468 | 3732.7251 | 59681.7275
Ex^2 | 4042088.5909 | 4042709.9033 | 4043181.5203 | 4043807.4137 | 4044023.1608 | 4045103.718 | 4046816.5897 | 4048139.9568 | 4048708.2536 | 4049279.7227 | 4050200.1184 | 4050993.8786 | 4051473.7423 | 4052472.2355 | 4052621.2761 | 4053057.6614 | 1036234584.0746
UNLOAD=b2 | -0.08779 | -0.17685 | -0.25761 | -0.27317 | -0.26338 | -0.34451 | -0.38143 | -0.29308 | -0.36892 | -0.4846 | -0.46554 | -0.48774 | -0.54325 | -0.5082 | -0.48427 | -0.47692 | -5.8973
| 392.2276 | 392.2168 | 392.2846 | 392.3695 | 392.4268 | 392.5442 | 392.5893 | 392.6932 | 392.7756 | 392.8732 | 392.9042 | 392.9804 | 392.9429 | 392.8305 | 392.9114 | 392.9232 | 6282.4935
| 784.3816 | 784.6162 | 784.6873 | 784.6904 | 784.7991 | 784.888 | 784.9905 | 785.2514 | 785.2279 | 785.3288 | 785.5219 | 785.616 | 785.6182 | 785.6365 | 785.7435 | 785.7784 | 12562.7758
| 1176.865 | 1177.025 | 1177.1935 | 1177.1633 | 1177.2633 | 1177.4155 | 1177.7 | 1177.9602 | 1177.9006 | 1178.0234 | 1178.1159 | 1178.2619 | 1178.3788 | 1178.4487 | 1178.3933 | 1178.4358 | 18844.5443
| 1373.5029 | 1373.5408 | 1373.5493 | 1373.4934 | 1373.5507 | 1373.6954 | 1374.0843 | 1374.2478 | 1374.452 | 1374.5133 | 1374.5528 | 1374.6254 | 1374.7584 | 1374.7907 | 1374.9524 | 1374.9797 | 21987.2892
AVERAGE | 745.3779 | 745.4444 | 745.4914 | 745.4887 | 745.5553 | 745.6397 | 745.7965 | 745.9719 | 745.9974 | 746.0508 | 746.1259 | 746.1992 | 746.231 | 746.2396 | 746.3033 | 746.328 | 11934.2411
Ex | 3726.8893 | 3727.2219 | 3727.4571 | 3727.4434 | 3727.7766 | 3728.1987 | 3728.9827 | 3729.8595 | 3729.9872 | 3730.254 | 3730.6293 | 3730.9959 | 3731.155 | 3731.1982 | 3731.5164 | 3731.6402 | 59671.2055
Ex^2 | 4040618.4574 | 4041458.7612 | 4042043.768 | 4041890.545 | 4042499.0449 | 4043486.6332 | 4045421.6046 | 4046975.1339 | 4047423.739 | 4048116.7775 | 4048771.0703 | 4049522.367 | 4050137.5082 | 4050331.5812 | 4050877.3655 | 4051116.5522 | 1035850831.1131
AVERAGE SUM | 745.3939 | 745.4686 | 745.5115 | 745.5537 | 745.6022 | 745.7024 | 745.8544 | 746.0288 | 746.0426 | 746.1207 | 746.2185 | 746.2848 | 746.3006 | 746.3678 | 746.4063 | 746.4365 | 11935.2933
Sum EX | 7453.9386 | 7454.6856 | 7455.1145 | 7455.5367 | 7456.022 | 7457.0241 | 7458.5442 | 7460.2882 | 7460.4262 | 7461.2071 | 7462.1851 | 7462.848 | 7463.0061 | 7463.6782 | 7464.0632 | 7464.3653 | 119352.933
Sum Ex^2 | 8082707.0483 | 8084168.6645 | 8085225.2884 | 8085697.9586 | 8086522.2058 | 8088590.3512 | 8092238.1943 | 8095115.0906 | 8096131.9926 | 8097396.5002 | 8098971.1887 | 8100516.2456 | 8101611.2505 | 8102803.8167 | 8103498.6416 | 8104174.2136 | 2072085415.1877
(SUM_EX)^2 | 55561200.7254 | 55572337.5109 | 55578732.5239 | 55585026.8108 | 55592264.1505 | 55607208.8964 | 55629880.9965 | 55655900.2527 | 55657959.3646 | 55669611.324 | 55684205.762 | 55694100.4463 | 55696460.0063 | 55706492.1953 | 55712239.4928 | 55716748.6249
| A1 | A2 | A3 | A4 | A5 | A6 | A7 | A8 | A9 | A10 | A11 | A12 | A13 | A14 | A15 | A16 | ES^2
| -0.47695 | -0.49942 | -0.54381 | -0.49804 | -0.49687 | -0.61992 | -0.77868 | -0.53832 | -0.9607 | -0.85581 | -0.67118 | -0.74889 | -1.1359 | -0.94055 | -0.85762 | -0.92322 | 133.3084
| 784.0637 | 784.2293 | 784.2683 | 784.4509 | 784.5339 | 784.8217 | 784.9615 | 785.3239 | 785.5086 | 785.7051 | 785.7974 | 785.8795 | 785.9359 | 785.8837 | 786.0007 | 786.0545 | 157839484.8595
| 1569.1556 | 1569.3281 | 1569.378 | 1569.5548 | 1569.7889 | 1570.0554 | 1570.4924 | 1570.7928 | 1570.8957 | 1571.0026 | 1571.2993 | 1571.4699 | 1571.684 | 1571.7943 | 1571.6646 | 1571.9001 | 631529796.2354
| 2353.9937 | 2354.309 | 2354.5847 | 2354.5712 | 2354.7429 | 2355.1112 | 2355.5604 | 2356.0454 | 2356.1477 | 2356.3558 | 2356.5477 | 2356.8769 | 2356.982 | 2357.2131 | 2357.3626 | 2357.2416 | 1420810930.4466
| 2747.2025 | 2747.3186 | 2747.4273 | 2747.4578 | 2747.4532 | 2747.6558 | 2748.3085 | 2748.6644 | 2748.835 | 2748.9994 | 2749.2118 | 2749.3706 | 2749.5402 | 2749.7277 | 2749.8929 | 2750.0923 | 1933990433.4441
EAiSi | 16165413.5904 | 16168337.1337 | 16170450.3383 | 16171395.5201 | 16173044.1015 | 16177180.3997 | 16184476.0343 | 16190230.0468 | 16192263.6153 | 16194792.7706 | 16197942.1335 | 16201032.2375 | 16203222.2448 | 16205607.1853 | 16206996.8759 | 16208348.1106
Note on H6: Alejandro: Siempre corresponde con grupos independientes
Note on A8: Alejandro: Siempre corresponde con grupos independientes